IS can script

Please translate, put bold text where stated and spaces.

| 1 | The Intersurgical IS Can Medical grade CO2 absorbers |  |
| --- | --- | --- |
| 2 | For many years Intersurgical has been producing medical grade co2 absorbents for use within anaesthetic breathing systems. |  |
| 3 | Amongst our range of absorbents, is the pre-filled IS Can absorber, which has been designed specifically to be used with GE machines. |  |
| 4 | The IS Can may be used with the following GE anaesthetic machines: Aisys, Avance and Aespire. |  |
| 5 | The IS Can may be fitted onto these machines either directly or via the CO2 bypass. The IS Can may also be used with the GE ADU carestation. |  |
| 6 | Fast and easy exchange The IS Can enables fast exchanges of absorbent without the need for time consuming pouring of loose granules, or the risk of contamination. |  |
| 7 | No dust emission Foam filters ensure that any fine particles or dust are retained within the IS Can and not allowed to escape into the breathing system or into the atmosphere during handling. |  |
| 8 | Efficient CO2, absorption A gas dispersion chamber at the base of the IS Can allows for optimum flow up through the absorbent. |  |
| 9 | The IS Can is available with a choice of two medical grade absorbents: Spherasorb Medical grade soda lime White to violet Loflosorb Alkali-hydroxide-free absorbent Green to violet |  |
| 10 | The IS Can is also available with spherasorb medical grade soda lime Pink to White. |  |
| 11 | Connection via the C02 bypass enables replacement of the IS Can without loss of ventilation. However, CO2 is not being removed when the canister is not engaged. |  |
| 12 | Before use, remover twin seal cap. |  |
|  | Before fitting ensure that the granules inside the IS CAN are settled and level. This is to provide optimum performance. This can be achieved by a few gentle taps. |  |
| 14 | With the handle facing the front of the machine: slide the IS Can into the cradle. |  |
| 15 | Push gently up until the blue clip engages over the front teeth of the IS Can. |  |
| 16 | The IS Can is now correctly fitted and ready for the pre-use leak test. |  |
| 17 | To remove the IS Can press the blue lever to allow the cradle to drop down and slide out the IS Can. |  |
| 18 | Connecting the IS Can directly to the GE machines. |  |
| 19 | With the handle facing the front of the machine: Hook the back feet of the IS Can around the pins at the front of the fitting station. |  |
| 20 | Push the IS Can gently up until the blue clip engages over the front feet of the IS Can. |  |
| 21 | The IS Can is now correctly fitted and ready for the pre-use leak test. |  |
| 22 | To remove press the blue lever and disengage the IS Can from the pins. |  |
| 23 | The IS Can may also be used with the GE ADU carestation. |  |
| 24 | Location pins. |  |
| 25 | Remove twin seal caps. With the IS Can in the upside down position and the handle away from the location pins: Hook the feet of the IS Can around the location pins of the ADU compact block. |  |
| 26 | Push the IS Can gently down until the clip engages over the opposite feet of the IS Can. |  |
| 28 | To remove, push down on the black lever and disengage the IS Can from the pins. |  |
| 29 | Aisys, Avance, Aespire, ADU and Compact are registered trade marks of GE Healthcare. |  |
|  | Always refer to the instructions for use provided in each box before using the IS CanTM |  |
|  | Quality, innovation and choice Interact with us |  |