|  | English | Your Language |
| --- | --- | --- |
| 01 | Most medical gas companies produce analgesic gas mixture that, although widely used, is known around the world by different brand names such as Kalinox Entonox and Livopan. |  |
| 02 | Regardless of the name this gas mixture is made up of 50% Oxygen and 50% Nitrous oxide and is mostly administered for pain relief both in and out of the hospital. |  |
| 03 | Typical areas of use would be Obstetrics, Trauma and as a inhaled analgesic for exploratory procedures such as Bronchoscopy. |  |
| 04 | Two of these methods use a demand valve which allows gas to flow to the patient, when the patient requires it. |  |
| 05 | The third method uses a free-flow device where the gas is flowing constantly. |  |
| 06 | Lets look at the demand valve methods in more detail as the key difference is in the use of a breathing system. |  |
| 07 | In the first variation of this method, the demand valve is connected to the top of a cylinder, whilst a breathing system is attached to the outlet port of the demand valve. |  |
| 08 | The patient holds the other end of the breathing system which may have either a face mask or mouth piece attached. |  |
| 09 | If the breathing system is to be for multi-patient use a suitable filter should be used for each patient. |  |
| 10 | When the patient needs to receive pain relief they will initiate a breath by drawing on the breathing system. |  |
| 11 | This creates a small negative pressure That triggers the demand valve to open allowing gas to flow. |  |
| 12 | When the patient starts to exhale the demand valve closes and the gas flow stops until the next inhalation. |  |
| 13 | The second demand valve method does not require a breathing system and is simply connected to the gas source by a length of pressure hose. |  |
| 14 | Normally a cylinder designed for administration of this gas and held by the patient. |  |
| 15 | On the outlook port of the demand valve a mouth piece with a filter or a face mask with a suitable filter is attached depending on patients choice. |  |
| 16 | This second variation works on the same mechanical principal as the first. |  |
| 17 | Both demand valve systems can also be fitted to a hospitals mains central gas supply via a wall outlet if available. |  |
| 18 | The third method of delivery Free-flow reverts back to the use of a breathing system connected to a flow metre on the wall outlet or to a specific gas cylinder that has a built in flow regulator. |  |
| 19 | The flow rate is set to the optimum level, typically 15 litres per minute for adults, slightly less for children. |  |
| 20 | This will fill the reservoir bag on the breathing system ready for the patient to take a breath. |  |
| 21 | The patient will either use a mouth piece or a face mask to breath the gas mixture through the breathing system. |  |
| 22 | When the patient takes a breath there would normally be sufficient volume in the reservoir bag, but if more is needed then the high constant flow of gas will make up the difference. |  |
| 23 | The reservoir bag is then rapidly filled again in readiness for the next patient breath. |  |
| 24 | Intersurgical has the right product for all these delivery systems applicable to this analgesic gas mixture of 50 % oxygen and 50% Nitrous oxide. |  |
|  | Titles on screen only. No Narration | Not for voice over artist |
|  | O2/N2O Systems |  |
|  | Kalinox Entinox Livopan |  |
|  | 50% Oxygen 50% Nitrous oxide |  |
|  | Mostly administered for pain relief. |  |
|  | Aneas of use Obstetrics Trauma Inhaled analgesic |  |
|  | Demand valve systems |  |
|  | Free-flow system |  |
|  | Demand valve |  |
|  | Outlet port |  |
|  | Clear-Guard 3 breathing filter |  |
|  | Intersurgical systems 2612000, 2612001, 2612010 |  |
|  | Flow rate Adults: 15L/min Children: Clinical judgement |  |
|  | Intersurgical. Quality, innovation and choice. |  |