| Fitting the Pyramid. V130306 |  |  |
| --- | --- | --- |
|  | English | (your language) |
| 01 | Intersurgical |  |
| 02 | Medical grade absorbents |  |
| 03 | Fitting the Pyramid™ |  |
| 04 | Specifically designed to be used with Draeger® anaesthetic machines which use the Draeger CLIC® system. |  |
| 05 | Compatible with Zeus®, Primus®, Perseus®, Fabius® CE, Fabius® GS and Julian®. Using either CLIC® MX00012 adaptor or CLIC® MX50090 adaptor |  |
| 07 | The Pyramid enables fast exchange of absorbent without the need for time consuming pouring of lose granules. |  |
| 08 | An air-tight cap protects the absorbent from moisture and the carbon dioxide in the air during storage. |  |
| 09 | Foam filters ensure that any fine particles or dust are retained and not allowed to escape into the breathing system. |  |
| 10 | A clip is provided to indicate the correct location when fitted to the Draeger MX50090 CLIC adaptor. |  |
| 11 | The gas dispersion chamber at the base of the Pyramid provides optimum flow and dispersion resulting in better utilization of the soda lime granules. |  |
| 12 | The Pyramid is available with a choice of two medical grade absorbents: |  |
| 13 | Spherasorb® Medical Grade Soda Lime White to violet |  |
| 14 | LoFloSorb® Alkali-hydroxide-free absorbent Green to violet |  |
| 15 | Using the Pyramid with the Draeger CLIC adaptor MX50090. |  |
| 16 | Remove the cap. |  |
| 17 | Ensure that the green clip is down and facing away from the adaptor. |  |
| 18 | Gently slide the Pyramid as far as it will go into the adaptor.. |  |
| 19 | Gently push the Pyramid to close the adaptor until there is a click. |  |
| 20 | Raise the green clip behind the adaptor button until there is a slight click. |  |
| 21 | The Pyramid is now correctly located within the adaptor and ready for the pre-use leak and compliance tests. |  |
| 22 | If the clip can not be raised behind the adaptor button, re-open the adaptor and reposition the Pyramid correctly. |  |
| 23 | Using the Pyramid with the Draeger CLIC adaptor MX00012. |  |
| 24 | Remove the cap. |  |
| 25 | Ensure that the green clip is down and facing away from the adaptor. |  |
| 26 | Gently slide the Pyramid as far as it will go into the adaptor. |  |
| 27 | Gently push the Pyramid to close the adaptor until there is a click. |  |
| 28 | Do not raise the green clip. It is not used with the CLIC adaptor MX00012. |  |
| 29 | The Pyramid is now correctly located within the adaptor and ready for the pre-use leak and compliance tests. |  |
| 30 | The new Pyramid can also be used with Draeger Infinity® empowered RFID anaesthetic machines. |  |
| 31 | The RFID function for the soda lime canister can be either permanently disabled, or overridden at the time of attachment of the Pyramid. |  |
| 32 | The Pyramid will function normally but without the RFID features. |  |
| 33 | Pyramid |  |
| 34 | Quality, innovation and choice |  |
| 35 | Draeger®, Zeus®, Primus®, Perseus®, Fabius® Julian®, CLIC® and Infinity® are trademarks of Draeger Medical. |  |