Auto-Fill Humidifier Chamber Translation document

|  | English | Your Language |
| --- | --- | --- |
| 1 | Dual float auto fill humidification chamber |  |
| 2 | 1. Inline filter 2. Strong clearly labelled material & water level indicator 3. Dual value / float 4. Fill set spike with hydrophobic membrane filter 5. Shrouded heated plate cover |  |
| 3 | 1 Inline filter Prevents any debris from entering the chamber. |  |
| 4 | 2 Strong clearly labelled material & water level indicator |  |
| 5 | The polycarbonate chamber provides easy visual assessment and durability. The yellow floating ring allows the user to ensure that the water level is between the clear markings. |  |
|  | 3 Dual float / valve |  |
| 4 | Primary float When the chamber fills to its MAX level, the primary float generates enough force to close the primary valve seal, stopping the flow of water into the chamber. |  |
| 6 | Secondary float Should the primary seal become compromised by a small particle, the secondary float is initiated. This generates a greater force which reshapes the valve onto the secondary seal, stopping the flow of water into the chamber. |  |
| 7 | Secondary Seal |  |
| 8 | 4 Fill set spike with hydrophobic membrane filter. |  |
| 9 | The sealed spike cover allows the user to conduct a ventilator leak test. The integral hydrophobic membrane filter allows for the use of rigid reservoir bottles. |  |
| 10 | 5 Shrouded heated plate cover |  |
| 11 | Prevents accidental burns when removing the chamber from the heater base. |  |
| 12 | Intersurgical |  |
| 13 | Quality, innovation and choice. |  |
|  | YouTube links |  |
| 14 | Workings of the Intersurgical Auto-Fill Humidification Chamber with dual float function for secondary back up security to prevent overfilling. |  |