7000 humidifier script

| 1 | Hi I'm Amy and I'm going to show you how to set up a 7000 series respiratory humidifier. |  |
| --- | --- | --- |
| 2 | You will be using a humidifier base, a humidifier chamber, a heated wire breathing system, a bag of sterile water, a temperature probe assembly, a heated wire adaptor, and breathing filters. |  |
| 3 | First put the humidifier into the ventilator bracket and connect to the mains. |  |
| 4 | Place the humidification chamber onto the humidifier and ensure it clicks into place. |  |
| 5 | Ensure the top of the humidifier chamber is placed lower than the patient airway position. |  |
| 6 | Add water to the humidifier chamber. In this case we are using an auto fill chamber. |  |
| 7 | To protect the ventilator we are using a low volume low resistance filter on the expiratory limb. |  |
| 8 | And a high efficiency filter on the inspiratory limb. |  |
| 9 | Remove the cover and attach the short inspiratory limb. |  |
| 10 | Connect the breathing system. The connector with the temperature port goes on the chamber. |  |
| 11 | The expiratory limb goes on the ventilator. |  |
| 12 | Now connect the temperature probe. |  |
| 13 | Connect he heated wire adaptor. |  |
| 14 | connect the clover leaf to the inspiratory limb |  |
| 15 | And the oval connector goes to the expiratory limb. |  |
| 16 | Connect the ventilator to the gas supply and turn on |  |
| 17 | Carry out the pressure testing applicable to the ventilator protocol. |  |
| 18 | Ensure there is enough water in the chamber and then turn on the humidifier. |  |
| 19 | Turn the humidifier on by briefly holding the on off button. |  |
| 20 | The unit will then self test displaying all lights. |  |
| 21 | Then it will display the software version. |  |
| 22 | Finally it will revert to the last settings used. |  |
| 23 | The humidifier has coloured lights for each of the probes. |  |
| 24 | When highlighted green it indicates which probe temperature is shown in the displace. This can be switched briefly by holding the 'Alarm - mute' button. |  |
| 25 | When highlighted yellow the probe is detecting a lower temperature than is set. |  |
| 26 | You can mute this alarm for two minutes by pressing the 'Alarm - Mute' button. |  |
| 27 | The alarm is disabled while the humidifier warms up. |  |
| 28 | When highlighted red the probe is detecting a higher temperature than is set. |  |
| 29 | You can mute this alarm for two minutes by pressing the 'Alarm - Mute' button. |  |
| 30 | Any alarm condition left unresolved for 10 minutes results in the humidifier entering the safety Stand-by Mode. The display will show '000'. |  |
| 31 | The unit must then be turned off and on to resume humidification. |  |
| 32 | Press and hold the MODE button to cycle through: Invasive mode - Non-Invasive mode - and Manual mode |  |
| 33 | On entering Manual mode the humidifier will first display the current set airway temperature - followed by the current set chamber temperature. |  |
| 34 | The airway temperature light will then flash indicating that you can now set your preferred temperature at the airway probe using the up and down buttons. |  |
| 35 | Holding the SET button allows the chamber temp to be set using the up and down buttons. |  |
| 36 | As a safety feature the option to set the temperature is limited to 5 seconds of inactivity. |  |
| 38 | To re-adjust the set temperatures simply hold down the up or down button for two seconds. |  |
| 39 | That's the basic set-up. Please refer to the instructions for use for more information |  |
| 40 | Thanks for watching. |  |

| On Screen titles |  |  |
| --- | --- | --- |
|  | Setting up a 7000 series humidifier |  |
|  | Pick list |  |
|  | Humidifier base |  |
|  | Humidifier chamber |  |
|  | Heated wire breathing system |  |
|  | Temperature probe assembly |  |
|  | Heated wire adaptor |  |
|  | Breathing filters |  |
|  | Ensure the top of the humidifier chamber is placed lower than the patient airway entry position. |  |
|  | Ensure markers are aligned when connecting the leads. |  |
|  | Connect the ventilator to the gas supply and turn on. |  |
|  | Carry out pressure testing applicable to the ventilator protocol. |  |
|  | Turn on humidifier |  |
|  | Press and hold the on/off button. |  |
|  | Unit will self test displaying all lights. |  |
|  | Unit will display the software version. |  |
|  | Unit reverts to the last settings used. |  |
|  | Coloured lights for each temperature probe |  |
|  | Patient end probe |  |
|  | Chamber probe |  |
|  | Green light |  |
|  | Indicates which probe temperature is displayed. |  |
|  | Holding the ALM button will display the other probe temperature. |  |
|  | yellow light |  |
|  | Indicates that the probe is detecting a lower temperature than is set. |  |
|  | You can mute the alarm for 2 minutes by pressing ... (note for translator there is an image of the button that should be pressed) |  |
|  | This alarm is disabled while the humidifier warms up. |  |
|  | Red light |  |
|  | Indicates that the probe is detecting a higher temperature than is set. |  |
|  | You can mute the alarm for 2 minutes by pressing ... (note for translator there is an image of the button that should be pressed) |  |
|  | Any alarm condition left unresolved for 10 minutes results in the humidifier entering the Safety Stand-by Mode. The display will show '000'. |  |
|  | The unit must then be turned off and on to resume humidification. |  |
|  | Modes |  |
|  | Invasive mode Patient end temp 40°C Chamber end temp 37°C |  |
|  | Non-Invasive mode Patient end temp 34°C Chamber end temp 31°C |  |
|  | Manual mode. |  |
|  | On entering manual mode humidifier will first display the current set airway temperature. |  |
|  | It will then display the current set chamber temperature. |  |
|  | The airway temp light will flash. |  |
|  | The airway temperature can now be set using the up and down buttons. |  |
|  | Holding the SET button allows the chamber temperature to be set using the up and down buttons. |  |
|  | For safety the option to set the temperature is limited to 5 seconds of inactivity. |  |
|  | To re-adjust the set temperatures simply hold down the up or down button for two seconds. |  |