CaStar R NEXT

For NIV

| 1 | This video shows the process of setting up, fitting and removing the CaStar R NEXT hood. Always refer to the Instructions for Use provided with each product. This video is not intended to replace the Instructions for Use. |  |
| --- | --- | --- |
| 2 | CaStar R NEXT For NIV Fitting and removing |  |
| 3 | Hood without underarm straps for non-invasive mechanical ventilation (NIV) Available in six sizes to comprehensively cover the broadest range of adult patients |  |
| 4 | Size information is on the top and side of each box. |  |
| 5 | A tape measure is included on the outside so that the correct size can be selected without opening the box. |  |
| 6 | Inside the box you will find ... Castar R Next hood Rigid annular body Inflation bulb Instructions for use Ear Plugs |  |
| 7 | Earplugs are included if you wish to use them. |  |
| 8 | Setting up the hood. |  |
| 9 | Open the bag using the perforated edge. |  |
| 10 | Reshape the hood. |  |
| 11 | Remove the patient access port cap. |  |
| 12 | Check the smooth operation of the anti-asphyxiation valve and place to one side. |  |
| 13 | Attach 22mm breathing system inspiratory and expiratory limbs. |  |
| 14 | Start the ventilator. |  |
| 15 | Positioning and activating the hood |  |
| 16 | Two people firmly open the collar using their finger tips and thumbs along the rigid ring, so that the patient's head can pass through. |  |
| 17 | Position the inflatable cushion around the patient's neck and hook the fasteners onto the two front metal connectors. |  |
| 18 | Take the rigid annular body, open it and position it around the patient’s neck following the indications. |  |
| 19 | Close the rigid annular body behind the patient's neck. |  |
| 20 | Ensure that the 'FRONT' arrows are aligned. |  |
| 21 | Using the four fastening tabs hook the rigid annular body onto the metal connectors positioned on the ring. |  |
| 22 | Inflate the cushion with the inflation bulb. The cushion should be full, but not under pressure. |  |
| 23 | Then close the clamp. |  |
| 24 | Fit the patient access port cap. If required pull the anti-asphyxiation button until the hood inflates. |  |
| 25 | You may remove the pump if you feel no further adjustment is necessary. |  |
| 26 | Push the inflation tube into the slot under the rigid annular body. |  |
| 27 | The hood is now set up correctly. |  |
| 28 | Using the patient access port |  |
| 29 | Twist the patient access port anticlockwise to access the patient’s face. |  |
| 30 | Lock the patient access port by twisting it clockwise. |  |
| 31 | If required pull the valve button using two fingers until the hood is pressurised. |  |
| 32 | Using sealed accesses for probes or catheters |  |
| 33 | Remove the access, position the probe/catheter or NG tube, then refit the access. |  |
| 34 | Removing the hood |  |
| 35 | Open the patient access port. |  |
| 36 | Probes or catheters that remain attached to the patient should be fed back into the hood. |  |
| 38 | Remove the rigid annular body. |  |
| 39 | Move the inflated cushion to the back of the patient's neck. |  |
| 40 | Two people firmly open out the collar and remove the hood. |  |
| 41 | Then turn off the ventilation flow. |  |
| 42 | CaStar R NEXT For NIV |  |
| Title | StarMed CaStar R Next Hood for NIV therapy |  |
| Description | The CaStar - R Next Hood for NIV therapy is a comfortable, versatile and lightweight patient interface designed to offer an alternative to traditional methods of non-invasive ventilation. |  |
| Tags | NIV 'maskless' 'mask less' 'respiratory care' 'non-invasive ventilation' |  |