DuoFlow How To Video Script.

|  | English |  |
| --- | --- | --- |
| Title | How to set up a DuoFlow™ single limb bi-lumen breathing system. |  |
| 01 | This video is a brief introduction to the DuoFlow™ bi lumen breathing system. The DuoFlow bi-lumen breathing system must always be used in accordance with the Instructions for Use supplied with the breathing system. This video is not a replacement for the Instructions for Use. |  |
| 02 | DuoFlow™ Breathing system |  |
| 03 | DuoFlow Inspiratory fresh gas Expiratory gas Improved patient safety Improved thermal efficiency Reduced plastic waste |  |
| 04 | Setting up DuoFlow in comparison to a traditional circle system. |  |
| 05 | Connect the tapered connector of the T-piece to the tapered connector of the inspiratory fresh gas flow port, using a push and twist action. |  |
| 06 | Then connect the short length of tubing to the expiratory port using a push and twist action. |  |
| 07 | Pre-use checks Checking the integrity of the inner lumen. |  |
| 08 | Disconnect the inspiratory limb and connect the expiratory limb to a gas source. |  |
| 09 | Disconnect the DuoFlow patient connection from the machine pressure test port and elbow or catheter mount. |  |
| 10 | Insert the red inner lumen test tool into the patient connection. |  |
| 11 | Pressurise the expiratory limb of the breathing system to 30cmH2O and ensure there is no decrease in pressure. This confirms the integrity of the inner lumen of the breathing system. |  |
| 12 | Re-configure the breathing system onto the anaesthetic machine ready for use. |  |
| 13 | Re-attach the elbow or catheter mount and return it to the leak testing port. |  |
| 14 | Return the tapered connector of the expiratory limb to the expiratory port, using a push and twist action. |  |
| 15 | Return the tapered connector of the T-piece to the inspiratory fresh gas flow port, using a push and twist action. |  |
| 16 | Attach the monitoring line between the anaesthetic work station and luer lock spigot. |  |
| 17 | Complete pre-use checks as specified by manufacturer of the anaesthetic work station. |  |
| 18 | Once a leak test of the full system has been performed, confirming no leaks or occlusions, connect the breathing system to the patient. |  |
| 19 | DuoFlow™ Breathing system |  |
|  | Intersurgical |  |
|  | Quality, innovation and choice |  |
|  | Interact with us |  |
|  | www.intersurgical.com |  |