| 01 | WARNING: This set-up and how to use video is not intended to replace the device Intructions For Use (IFU), read and familiarise yourself with the device IFU before first use. |  |
| --- | --- | --- |
| 02 | InterActive™ Respiratory humidifier Set-up and how to use |  |
| 03 | The InterActive respiratory humidifier provides a new solution for active humidification and combines advanced humidifier technology with Intersurgical’s comprehensive range of heated wire breathing systems and accessories. |  |
| 04 | Top mounted leads Colour coded for quick and easy set up |  |
| 05 | Alarm indication Pictorial alarm display for fast error identification and resolution |  |
| 06 | Touch-screen display The intuitive user interface is navigated via a 4.2” colour touch-screen |  |
| 07 | Mode selection Invasive, non-invasive and manual mode options, including last used settings memory |  |
| 08 | IR water level sensing Immediate alarm notification before humidification is compromised |  |
| 09 | You will need |  |
| 10 | Heated wire adaptor lead (included) |  |
| 11 | InterActive™ |  |
| 12 | Temperature probe (included) |  |
| 13 | Heated wire breathing system with humidification chamber |  |
| 14 | Setting up |  |
| 15 | Slide the humidifier on to an appropriate bracket via the mounting feature. |  |
| 16 | Push down on the InterActive housing grip feature. |  |
| 17 | Applying a slight downward pressure, slide the humidification chamber fully on to the top plate. |  |
| 18 | Unwind the fill set and remove the red cap from the spike. |  |
| 19 | Pierce the fluid reservoir with the fill set spike and hang the reservoir on the pole above the humidification chamber. |  |
| 20 | Ensure that the feed line is not kinked and that the humidification chamber begins to fill with water. |  |
| 21 | Remove the cassette from the humidification chamber ports. |  |
| 22 | Attach one end of the dry limb to either of the two chamber ports and the other end to the gas supply. |  |
| 23 | Connect the elbow on the inspiratory limb to the other port on the chamber. |  |
| 24 | Insert the yellow heated wire adapter connector into the yellow socket on the top of the humidifier base. |  |
| 25 | Connect the ‘clover-leaf’ end to the socket on the breathing system elbow above chamber. |  |
| 26 | Insert the blue temperature probe connector into the blue socket on top InterActive. |  |
| 27 | Insert the distal (short wire) temperature probe into the port on the inspiratory limb elbow at the chamber. |  |
| 28 | Push and twist |  |
| 29 | Insert the proximal (long wire) temperature probe into the port at the patient end of the inspiratory limb. |  |
| 30 | Connect the InterActive plug plug to the mains supply. |  |
| 31 | The splash screen will be displayed. A system test is initiated automatically and the InterActive alarms will beep twice. |  |
| 32 | Ensure flow of gas to the breathing system then to start humidification, press the power on/off. |  |
| 33 | Changing mode |  |
| 34 | Press the mode symbol on the therapy display screen to change the selected therapy mode. |  |
| 35 | Use the navigation arrows to cycle through the therapy mode options. |  |
| 36 | Press the  or the central mode symbol to select the mode. |  |
| 37 | Using the manual mode |  |
| 38 | Press the mode symbol. |  |
| 39 | Use the navigation arrows to cycle through to the gas temperature icon. Press the  or the central mode symbol to select. |  |
| 40 | To view or change the Manual mode target temperatures, press the current gas temperature. |  |
| 41 | Select either the chamber or patient. |  |
| 42 | Press + or - to adjust,  to confirm. |  |
| 43 | Press the back arrow to return to the therapy screen. |  |
| 44 | Manual mode lock |  |
| 45 | Press the mode symbol on the therapy display screen to change the selected therapy mode. |  |
| 46 | Navigate to the settings cog and select. |  |
| 47 | The default factory setting is that temperature modification is un-locked. |  |
| 48 | Press the temperature padlock icon to toggle the set temperature permissions in manual mode. |  |
| 49 | The user is now unable to edit the set temperatures in the manual mode without unlocking first. |  |
| 50 | Turning off |  |
| 51 | Press the power on/off symbol. |  |
| 52 | Press  to confirm. |  |
| 53 | Intersurgical |  |
| 54 | Quality, innovation and trust |  |
| 55 | Interact with us |  |
| 56 | intersurgical.com |  |