InterCooler set up script

| 1 | InterCooler™ Respiratory dehumidifier |  |  |
| --- | --- | --- | --- |
| 2 | How to add the InterCooler to a dual heated wire breathing system |  |  |
| 3 | Remove the humidifier from the ventilator bracket. |  |  |
| 4 | Mount the humidifier on to the tandem InterCooler bracket. |  |  |
| 5 | Mount the InterCooler on to the tandem InterCooler bracket. |  |  |
| 6 | Setting up |  |  |
| 7 | Insert the fluid collection bag into the InterCooler. |  |  |
| 8 | Insert the InterCooler cassette, ensuring that the rear edge is located beneath the lip of the InterCooler base. |  |  |
| 9 | Push the cassette down and the latch forward, ensure location by checking the latch is flush with the front of the InterCooler housing. |  |  |
| 10 | Connect the expiratory heated wire limb to the 22m male connector of the InterCooler cassette. |  |  |
| 11 | Connect the return limb to the 22m female cassette connector. |  |  |
| 12 | If filtration is required use the 7 day Air-Guard Clear 7 filter. |  |  |
| 13 | Use the push and twist technique. |  |  |
| 14 | Turn on the InterCooler by pressing the power button. The button will illuminate green if the system is set up and fuctioning correctly. |  |  |
| 15 | How to change the fluid collection bag |  |  |
| 16 | When the fluid collection bag is full the power button will flash blue. |  |  |
| 17 | To replace the fluid collection bag press the InterCooler latch to release the cassette. |  |  |
| 18 | Remove the full collection bag and replace with a new bag. |  |  |
| 19 | Please note: Ventilation can continue uninterrupted during bag replacement. |  |  |
| 20 | Remember: Ensure location by checking that the latch is flush with the front of the InterCooler housing. |  |  |
| 21 | The button will illuminate green if the system is set up and fuctioning correctly. |  |  |
| 22 | InterCooler™ Reduce intervention, increase protection |  |  |
| 23 | Intersurgical Quality, innovation and choice Interact with us |  |  |