Intersurgical i-flow interFlow™

|  | English |  |
| --- | --- | --- |
| 1 | WARNING: This set-up video is not intended to replace the device Instructions For Use (IFU). Read and familiarise yourself with the device, and all associated device IFUs before first use. |  |
| 2 | How to set-up the InterFlow and i-floTM for HFO |  |
| 3 | Display Power button Proximal pressure port O2 supply connection Respiratory gas outlet Air supply connection |  |
| 4 | Touch screen Fast responsive screen |  |
| 5 | You will need the following equipment i-flo cannula Breathing system Humidifier Air and oxygen hoses Equipment stand InterFlow |  |
| 6 | Interflow set up |  |
| 7 | Connect the air and oxygen hoses to InterFlow’s gas inlets. Connect the hoses to the air and oxygen wall outlets. Connect to the mains power supply. |  |
| 8 | i-flo breathing system set-up |  |
| 9 | Attach the elbow using a push and twist action. If required, attach optional breathing filter using a push and twist action |  |
| 10 | Connecting the VHB20 humidifier |  |
| 11 | Place the humidification chamber on to the humidifier and ensure that it clicks into place. Add water to the humidification chamber. In this case we are using a auto fill chamber. |  |
| 12 | Attach the short inspiratory limb. Connect the breathing system. The connector with the temperature port goes on the chamber. |  |
| 13 | Connect the red heated wire adaptor. Connect the clover leaf to the inspiratory limb. |  |
| 14 | Connect the blue temperature probe adaptor and both probes. |  |
| 15 | 'Set-up HFOT therapy' |  |
| 16 | Device Start-up To switch on the device, press and hold the power on/off button. The splash screen will be displayed, a system test is initiated and the alarms will beep twice. |  |
| 17 | Therapy Screen Press the mode symbol on the therapy display screen to change the selected therapy mode. Use the navigation arrows to cycle through the therapy mode options. |  |
| 18 | Press the tick or the central mode symbol to select the mode. Once a new mode is selected, the mode confirmation screen is entered. |  |
| 19 | The current mode is shown on the left and the new mode on the right. Press the tick to confirm new mode symbol selection. |  |
| 20 | Treatment To start treatment, press the start treatment symbol on the therapy screen. To change the flow rate, press the parameter value on the therapy screen. |  |
| 21 | Therapy Parameter Adjustment Press – or + or tap the gauge to adjust the value. Press the tick to confirm selection and return to the therapy screen. |  |
| 22 | Repeat the same process to adjust the O2 percentage. To start treatment, press the start treatment symbol on the therapy screen and again to confirm. |  |
| 23 | Therapy Graph Display To change the graph information displayed, press the graph on the therapy screen to cycle. |  |
| 24 | Start humidifier Switch on the humidifier and select the appropriate starting temperature settings in line with hospital policy. |  |
| 25 | Fitting i-flo and starting HFOT |  |
| 26 | i-flo sizing and fitting Hold the guide up to the patient's nares to determine the correct size cannula. The cannual prongs should not create a seal with the nares or press on the septum. |  |
| 27 | Select the correct size cannula. Holding the strap away from the face, insert the nasal prongs into the patient's nares and lift the strap over their head. |  |
| 28 | Adjust the tension to ensure a comfortable fit. The 180° rotating elbow can be moved left or right for convenient positioning. |  |
| 29 | Attach the breathing system to the i-flo swivel connector using a push and twist action. |  |
| 30 | Simple. Intuitive. Advanced |  |
| 31 | Intersurgical Quality, innovation and Choice Interact with us www.intersurgical.com |  |