Intersurgical n-flow interFlow™

|  | English |  |
| --- | --- | --- |
| 1 | WARNING: This set-up video is not intended to replace the device Instructions For Use (IFU). Read and familiarise yourself with the device, and all associated device IFUs before first use. |  |
| 2 | How to set-up the InterFlow and nFlowTM for nCPAP |  |
| 3 | Display Power button Proximal pressure port O2 supply connection Respiratory gas outlet Air supply connection |  |
| 4 | Touch screen Fast responsive screen |  |
| 5 | You will need the following equipment InterFlow Air and oxygen hoses Humidifier Equipment stand nFlow breathing system Humidification chamber nFlow masks and prongs nFlow Cosyfit bonnet |  |
| 6 | Interflow set up |  |
| 7 | Connect the air and oxygen hoses to InterFlow’s gas inlets. Connect the hoses to the air and oxygen wall outlets. Connect to the mains power supply. |  |
| 8 | nFlow breathing system set-up |  |
| 9 | Attach the short inspiratory limb using a push and twist action. Attach the monitoring line using a push and twist action. |  |
| 10 | Set-up the nFlow breathing system with the VHB20 humidifier See ‘Set-up of the VHB20 humidifier’ for futher details @intersurgical |  |
| 11 | Size the nFlow interfaces and bonnet |  |
| 12 | nFlow CosyFit Nasal Mask |  |
| 13 | Hold the interface size guide up to the baby's nose to determine the best match nasal mask size. |  |
| 14 | Select the size that best fits the baby, ensuring the triangular outline exceeds the perimeter of the nose. |  |
| 15 | Nasal prong Hold the interface size guide up to the baby's nose to determine the best match nasal prong size. |  |
| 16 | Correct Orientation Connect the selected patient interface to the nFlow CPAP device ensuring the correct orientation. |  |
| 17 | CosyFit bonnet Determine the correct CosyFit bonnet size using the head size guide. |  |
| 18 | Select the appropriate bonnet size. |  |
| 19 | Fit the bonnet to the patient, ensuring the correct orientation of attachment pads. Tie the top of the bonnet if required. |  |
| 20 | Set-up nCPAP therapy (Pressure controlled mode) |  |
| 21 | To switch on the device, press and hold the power on/off button. The splash screen will be displayed, system test is initiated automatically and the alarms will beep twice. |  |
| 22 | Set the pressure to 5cmH2O. Press the start treatment symbol to start therapy. |  |
| 23 | Occlude the interface to confirm system performance. |  |
| 24 | Ensure that the pressure level is achieved and constant. Ensure that the pressure drops when prongs/mask are open. |  |
| 25 | To change the therapy parameter setting, press the parameter value on the therapy screen. |  |
| 26 | To change the graph information displayed, press the graph on the therapy screen to cycle. |  |
| 27 | Switch on the humidifier and select the appropriate starting temperature settings in line with hospital policy. |  |
| 28 | Fitting nFlow and starting nCPAP |  |
| 29 | nFlow CosyFit Adjust the distance from the patient interface (prong or mask) to the tube support to match the distance from the patient’s nares to forehead. |  |
| 30 | Check delivered pressure by occluding the patient interface and waiting until stable. |  |
| 31 | Position the patient interface over / into the nares and ensure correct alignment. |  |
| 32 | Whilst holding the nFlow CPAP device in position bring the tube support into contact with bonnet attachment pad. |  |
| 33 | Gently attach the straps to the side attachment pads using lateral positioning to ensure optimal patient comfort and ensure a good interface seal with minimal leaks. |  |
| 34 | Final adjustment of patient interface can be achieved by compressing and extending the exhaust tube and adjusting the angle of the patient interface. |  |
| 35 | CAUTION Overtightening the side straps may result in patient trauma, consider a different interface size or type if achieving good seal is difficult. |  |
| 36 | Ensure nCPAP therapy is being administered as intended. |  |
| 37 | InterFlow Simple. Intuitive. Advanced |  |
| 38 | Intersurgical Quality, innovation and Choice Interact with us www.intersurgical.com |  |