|  | English |  |
| --- | --- | --- |
| 01 | This video shows the process of setting up, fitting and removing a CaStar CPAP hood. Always refer to the Instructions for Use provided with each product. This video is not intended to replace the Instructions for Use. |  |
| 02 | CaStar CPAP Hood without manometer Fitting and removing |  |
| 03 | A patient interface for CPAP that is comfortable, versatile and lightweight. Available in six sizes to comprehensively cover the broadest range of adult patients. |  |
| 04 | CaStar CPAP hoods are available with a built-in manometer. |  |
| 05 | CaStar CPAP hoods are also available without a built-in manometer. |  |
| 06 | In this video we are using the version without a fixed manometer. |  |
| 07 | The CP929 (adaptor with built-in manometer) can be used to add manometer functionality if desired. |  |
| 08 | Push & Twist |  |
| 09 | Sizing |  |
| 10 | A tape measure can be found in the outer box. |  |
| 11 | Measure the patient's neck. |  |
| 12 | Size information is on the side of each box. |  |
| 13 | Inside the box you will find |  |
| 14 | CaStar CPAP Hood |  |
| 15 | Under arm straps |  |
| 16 | Instructions for use |  |
| 17 | Pre-use check of the hood |  |
| 18 | Remove the patient access port cap. |  |
| 19 | Check the smooth operation of the anti-asphyxiation valve and place to one side. |  |
| 20 | Attach the arm straps to the rear of the hood to suit the patient. |  |
| 21 | Setting up the hood for use on a flow driver |  |
| 22 | You will need |  |
| 23 | 1. CaStar CPAP hood 2. Breathing system 3. Filter 4. PEEP valve |  |
| 24 | Connect PEEP valve and filter to the expiratory side of the hood. |  |
| 25 | When treating infected patients, it is recommended to use a viral/bacterial filter in this position to prevent environmental cross contamination. |  |
| 26 | Set the PEEP valve to the desired pressure. |  |
| 27 | Attach a 22mm breathing system inspiratory limb using a push and twist action. |  |
| 28 | Optional: A breathing filter can be fitted between the hood and inspiratory limb to reduce noise for the patient. Use breathing filters according to hospital infection control protocols |  |
| 29 | Start the flow driver. |  |
| 30 | Fitting the hood |  |
| 31 | Two people firmly open the collar using their fingertips, and thumbs along the rigid ring so that the patient's head can pass through. |  |
| 32 | Raise the patients arms to ensure that the straps are tucked well under for strap sizing and attach the front straps. |  |
| 33 | Fit the access port cap then pull the anti-asphyxiation button until the hood inflates. |  |
| 34 | If needed adjust the straps so that the rigid ring rests about 1cm away from the shoulders. |  |
| 35 | When positioned correctly, you may fold and pin the front strap ends to enhance patient visibility. |  |
| 36 | The hood is now set up correctly. |  |
|  | You can use the access port for temporary patient access if needed. |  |
| 37 | Removing the hood |  |
| 38 | Remove the access port cap. |  |
| 39 | Undo the front straps. |  |
| 40 | Two people firmly open out the collar and remove the hood. |  |
| 41 | Stop the flow driver. |  |
| 42 | Quality, innovation and choice |  |
| 43 | Interact with us |  |
| 44 | www.intersurgical.com |  |