CaStar R NIV: Translation document

| 01 | This video shows the process of setting up, fitting and removing a CaStar R hood. Always refer to the Instruction for Use provided with each product. This video is not intended to replace the Instructions for Use. |  |
| --- | --- | --- |
| 02 | CaStar R Hood For NIV Fitting and removing |  |
| 03 | A patient interface for non-invasive ventilation that is comfortable, versatile and lightweight. |  |
| 04 | Available in six sizes to comprehensively cover the broadest range of adults patients. |  |
| 05 | Sizing |  |
|  | A tape measure can be found in the outer box or outside of the product package in some configurations. |  |
| 06 | Measure the patient's neck. |  |
| 07 | Size information is on the side of each box. |  |
| 08 | Open the bag using the perforated edge. |  |
| 09 | Inside the box you will find... |  |
| 08 | CaStar R Hood |  |
| 09 | Under arm straps |  |
| 10 | Inflation bulb* *included with some configurations and can be ordered seperately |  |
| 11 | Instructions for use |  |
| 12 | Sealed access ports cap |  |
| 13 | Ear plugs |  |
| 14 | Pre-use check of the hood |  |
|  | Reshape the hood |  |
| 15 | Remove the patient access port cap. |  |
| 16 | Check the smooth operation of the anti-asphyxiation valve and place to one side. |  |
| 17 | Attach the arm straps to the rear of the helmet to suit the patient. |  |
| 18 | Attach 22mm breathing system inspiratory and expiratory limbs using a push and twist action. |  |
| 19 | Note: It makes no difference on which side of the hood the inspiratory and expiratory limbs are connected. |  |
| 20 | Optional: A breathing filter can be fitted between the hood and inspiratory limb to reduce noise for the patient. Use breathing filters according to hospital infection control protocols |  |
| 21 | Attach the breathing system limbs to the ventilator using a push and twist action. |  |
| 22 | Start the ventilator. |  |
|  | Fitting the hood |  |
| 23 | Two people firmly open the collar using their finger tips, and thumbs along the rigid ring so that the patient's head can pass through. |  |
| 24 | Raise the patients arms to ensure that the straps are tucked well under for strap sizing and attach the front straps. |  |
| 25 | Fit the access port cap then pull the anti-asphyxiation button until the hood inflates. |  |
| 26 | If needed adjust the straps so that the rigid ring rests about 1cm away from the shoulders. |  |
| 27 | When positioned correctly, you may fold and pin the front strap ends to enhance patient visibility. |  |
| 28 | Inflate the cushion. The cushion should be full but not under pressure. |  |
| 29 | When suitably filled clamp off the cushion. You may remove the pump. |  |
| 30 | The hood is now set up correctly. |  |
| 31 | You can use the access port for temporary patient access if needed. |  |
| 32 | Hood removal |  |
| 33 | Remove the access port cap |  |
| 34 | Removing the hood |  |
| 35 | Deflate the cushion by opening the clamp. |  |
| 36 | Remove the patient access port cap. |  |
| 37 | Undo the front straps. |  |
| 38 | Two people firmly open out the collar and remove the hood. |  |
| 39 | Stop the ventilator |  |
| 40 | Newly designed sealed access ports are located on the hood wall closer to the patient’s airway. |  |
| 40b | Using the sealed access ports |  |
| 40b | Newly designed sealed access ports are located on the hood wall closer to the patient’s airway. |  |
| 41 | The large diameter of the flexible connectors, makes them suitable for single or multiple lumen probes or catheters. |  |
| 42 | The yellow ports can be used for catheters ≥ 12 Ch. |  |
| 43 | Where a smaller seal is needed for catheters ≤ 10 Ch, you can fit the replacement green port provided. |  |
| 44 | When a catheter is not in use, first fit the small cap into the larger collar. |  |
| 45 | Then insert the access port into its seat. |  |