CaStar Up: Translation document

| 01 | This video shows the process of setting up, fitting and removing the CaStar Up CPAP hood. Always refer to the Instructions for Use provided with each product. This video is not intended to replace the Instructions for Use. |  |
| --- | --- | --- |
| 02 | CaStar Up CPAP Hood Fitting and removing |  |
| 03 | A patient interface for CPAP that is comfortable, versatile and lightweight Available in six sizes to comprehensively cover the broadest range of patients from paediatrics to adults |  |
| 04 | Size information is on the top and side of each box. |  |
| 05 | Some configurations include a tape measure on the outside so that the correct size can be selected without opening the box. |  |
| 06 | In the box you will find ... |  |
| 07 | CaStar Up CPAP hood |  |
| 08 | Under arm straps |  |
| 09 | Inflation bulb (with some configurations) |  |
| 10 | Instruction for use |  |
| 11 | Ear plugs (with some configurations) |  |
| 12 | Setting up the device |  |
| 13 | Open the bag using the perforated edge. |  |
| 14 | Reshape the hood. |  |
| 15 | Remove the patient access port cap. |  |
| 16 | Check the smooth operation of the anti-asphyxiation valve and place to one side. |  |
| 17 | .Separate the two parts of the device (upper and lower rings). |  |
| 18 | Lengthen the under arm straps using the velcro adjustment. |  |
| 19 | Secure the underarm straps to the rear of the lower ring. |  |
| 20 | Positioning and activating the device |  |
| 21 | Two people firmly open the collar using their finger tips and thumbs along the ring, so that the patient's head can pass through. |  |
| 22 | Raise the patient arms to ensure that the straps are tucked well under for strap sizing. |  |
| 23 | Complete positioning of the fastening system on the front of the lower ring. |  |
| 24 | Connect the 22mm breathing system to the inspiratory connector. |  |
| 25 | Choose a fixed value or an adjustable PEEP valve. |  |
| 26 | It is recommended to set a PEEP of at least 5cm/H2O. |  |
| 27 | Connect the PEEP valve to the expiratory connector. |  |
| 28 | Start the flow driver. |  |
| 29 | Attach the upper part of the helmet to the lower ring. |  |
| 30 | Lock into place using the levers. |  |
| 31 | Fit the patient access port. If required pull the anti-asphyxiation button until the hood inflates. |  |
| 32 | Adjust the length of the straps so that the ring rests about 1 centimetre away from the patient's shoulders. |  |
| 33 | Inflate the cushion if fitted. The cushion should be full but not under pressure. |  |
| 34 | Close the clamp. You may remove the pump. |  |
| 35 | Using the patient access port |  |
| 36 | To open the patient access port turn anti-clockwise. |  |
| 37 | Lock the patient access port by twisting it clockwise. |  |
| 38 | If required pull the valve button using two fingers until the hood is pressurised. |  |
| 39 | Opening the hood for nursing procedures |  |
| 40 | Open the patient access port. |  |
| 41 | One operator can easily remove the upper part of the hood and have full access to the patient. |  |
| 42 | Fit the patient access port. |  |
| 43 | If required pull the anti-asphixiation button until the hood inflates. |  |
| 44 | Using the sealed accesses for probes and catheters |  |
| 45 | Remove the access, position the probe/catheter or NG tube, then refit the access. |  |
| 46 | Removing the helmet |  |
| 47 | Open the clamp to deflate the cushion. |  |
| 48 | Remove the patient access port cap. |  |
| 49 | Separate the upper part of the hood from the lower using the levers located on the rigid rings. |  |
| 50 | Turn off the flow driver. |  |
| 51 | Remove any probes/catheters used during the therapy from the airtight accesses. |  |
| 52 | Detach the fastening systems. |  |
| 53 | Two people firmly open out the collar and remove the lower part of the hood. Probes and catheters can remain in place on the patient. |  |
| 54 | CaStar Up CPAP Hood |  |
| 55 | Intersurgical Quality, innovation and choice |  |