| 00:04 | English |  |
| --- | --- | --- |
| 01 | This video is a brief introduction to the Ventumask 30 Pro CPAP mask. Intersurgical products must always be used in accordance with the Instructions for Use supplied with the device. This video is not a replacement for the Instructions for Use. Intersurgical accepts no responsibility for the medical practices of those who directly or indirectly reference this video. |  |
| 02 | Ventumask 30 Pro CPAP mask with Venturi flow driver and adjustable PEEP valve |  |
| 03 | Ventumask 30 PRO is a unique design for CPAP therapy that is simply activated by connecting to an appropriate oxygen supply. |  |
| 04 | It can be used in a hospital environment or pre-hospital, for example in an ambulance. |  |
| 05 | Features |  |
| 06 | Ventumask 30 PRO is a complete and easy to use system where the oxygen supply tubes, Venturi flow driver and a PEEP valve are integral to the CPAP mask |  |
| 07 | The built in PEEP valve can provide patient pressures in a range of settings up to 20 cmH₂O |  |
| 08 | O₂% can be adjusted from 30% to 100% at each PEEP level |  |
| 09 | The user is guided by a performance table integrated into the product, for independent PEEP and O₂% settings |  |
| 10 | Mask Fitting |  |
| 11 | Every box contains 5 masks and a size guide for each mask. |  |
| 12 | Place the tip of the size guide below the bottom lip as shown. |  |
| 13 | Check the nasal bridge alignment to identify the correct size. |  |
| 14 | Available in 4 adult sizes. |  |
| 15 | Check that the cushion is inflated. |  |
| 16 | If the cushion is not sufficiently inflated, add more air using a syringe. |  |
| 17 | How to set up Ventumask |  |
| 18 | Example of therapy setting: O₂% = 30% PEEP = 7.5 cmH₂O |  |
| 19 | Refer to the performance table to check what flows are required for the selected therapy setting. |  |
| 20 | To deliver O₂% = 30% and PEEP = 7.5 cmH₂O |  |
| 21 | Ensure that the Venturi valve is clicked into the open position. |  |
| 21a | This will allow appropriate air entrainment |  |
| 22 | Set the PEEP valve to 7.5 cmH₂O. |  |
| 23 | Check that line B is occluded using the blue pin. |  |
| 24 | Connect blue line A to the flow meter ensuring that it is securely fitted. |  |
| 25 | Activate the oxygen flow according to the table. |  |
| 26 | Fit the mask using the harness provided. |  |
| 27 | Check that the harness is comfortable for the patient. |  |
| 28 | Re-adjust if necessary. |  |
| 29 | You may fold over straps if they are obstructing. |  |
| 30 | Check the presence of pressure within the mask using the integrated manometer. |  |
| 31 | Delivering O₂% greater than 30% |  |
| 32 | Example of therapy setting: O₂% = 50% PEEP = 7.5 cmH₂O |  |
| 33 | Refer to the performance table to check what flows are required for the selected therapy settings. |  |
| 34 | Ensure that the Venturi valve is clicked into the open position. |  |
| 34a | This will allow appropriate air entrainment |  |
| 35 | Set the PEEP valve to 7.5 cmH₂O. |  |
| 36 | Connect blue line A to the first oxygen source. A single oxygen flowmeter or one side of a dual oxygen flowmeter. |  |
| 37 | Additional oxygen flow is required in line B. |  |
| 38 | Remove the blue pin from yellow line B and connect to the second oxygen source. |  |
| 39 | Activate the oxygen flows according to the table. |  |
| 40 | Changing the PEEP setting during therapy |  |
| 41 | Example with PEEP increased from 7.5 to 10 cmH₂O. |  |
| 42 | Adjust the PEEP valve. |  |
| 43 | Adjust the oxygen flows according to the table. |  |
| 44 | How to set up Ventumask |  |
| 45 | Example of therapy setting: O₂% = 100% At any setting on the PEEP valve |  |
| 46 | Refer to the performance table to check what flows are required for the desired therapy settings. |  |
| 47 | Ensure that the Venturi valve is clicked into the closed position. |  |
| 47a | This will stop air entrainment |  |
|  | Set the PEEP valve to the desired pressure. |  |
| 49 | Connect blue line A to the first oxygen source. |  |
| 50 | Remove the blue pin from yellow B and connect to the second O₂ source. |  |
| 51 | Activate the oxygen flows according to the table. For delivering O₂% = 100% at least 40 L/min must be reached. |  |
| 52 | Using the nebulisation feature |  |
| 53 | Remove the cap located on the nebulisation port and store in a safe place. |  |
| 54 | Connect the angled connector (provided) to the nebulisation port. |  |
| 54a | Use a push and twist action |  |
| 55 | Connect the nebuliser (not provided) to the angled connector. Note: The nebuliser should be positioned vertically by rotating the connector. |  |
| 55a | Use a push and twist action |  |
| 56 | Use a nebuliser that delivers a particle size of 2-5µm. |  |
| 57 | When using the nebuliser connector refer to the yellow table for specific settings. |  |
| 58 | Connect the nebuliser to the power source according to the manufacturer’s instructions. |  |
| 57 | Activate the oxygen flow according to the table. |  |
| 57a | No oxygen flow required in yellow line B |  |
| 59 | Set the PEEP valve to 5 cmH₂O. |  |
| 60 | At the end of therapy, disconnect the nebuliser and the angled connector. |  |
| 61 | Reposition the cap on the nebulisation port using a push and twist action. |  |
| 62 | Using a flip-top cap for bronchoscopy / aspiration |  |
| 63 | Open the upper cap to use the bronchoscope through the sealed hole. |  |
| 64 | Open the lower cap to use the suction cannula. |  |
| 65 | The overpressure valve will automatically activate at any time to ensure maximum patient pressure will not exceed 25 cmH2O. |  |
| 66 | Removing the Ventumask |  |
| 67 | Undo the harness and remove the mask. |  |
| 68 | Quality, innovation and trust |  |
|  | Promo |  |
| 01 | Features |  |
| 02 | Ventumask 30 PRO is a complete and easy to use system where the oxygen supply tubes, Venturi flow driver and a PEEP valve are integral to the CPAP mask |  |
| 03 | The built in PEEP valve can provide patient pressures in a range of settings up to 20 cmH₂O |  |
| 04 | O₂% can be adjusted from 30% to 100% at each PEEP level |  |
| 05 | The user is guided by a performance table integrated into the product, for independent PEEP and O₂% settings. |  |
| 06 | What’s new in Ventumask 30 PRO |  |
| 07 | Integrated manometer |  |
| 08 | Nebuliser port |  |
| 09 | Bronchoscopy port |  |
| 10 | Suctioning port |  |
| 11 | Overpressure relief valve |  |
| 12 | New size guide |  |