StarMed Ventumask with

Venturi flow driver for CPAP therapy

| 0 | This video shows the process of setting up, fitting and removing the Ventumask. Always refer to the Instructions for Use provided with each product. This video is not intended to replace the Instructions for Use. |  |
| --- | --- | --- |
| 1 | Ventumask CPAP mask with Venturi flow driver and adjustable PEEP valve |  |
| 2 | Ventumask is a unique design for CPAP therapy that is simply activated by connecting to an appropriate oxygen supply. It can be used in a hospital environment or pre-hospital use, for example in an ambulance. |  |
| 3 | Features • Ventumask is a complete and easy to use system, where the oxygen supply tubes, a Venturi flow driver and a PEEP valve are integral to the CPAP mask. • The adjustable PEEP valve is designed exclusively for Ventumask and the administration to the patient is facilitated at some pre-settings: 5 - 7.5 - 10 - 12.5 cmH₂O. • O₂% can be adjusted from 40% to 100% at each PEEP level. • The user is guided by a performance table integral to the product, for independent PEEP and O₂% settings. |  |
| 4 | Size information can be found on the packet.= |  |
| 5 | Available in 3 adults sizes S/M L XL |  |
| 6 | Open the packaging. |  |
| 7 | Inside the packet you will find ... Ventumask Harness Instructions |  |
| 8 | Preparing the Ventumask |  |
| 9 | Check the cushion is inflated. |  |
| 10 | If the cushion is not sufficiently inflated, add more air using a syringe. |  |
| 11 | How to set up Ventumask Example of therapy settings: O₂% = 40% PEEP = 7,5 CmH₂O |  |
| 12 | Refer to the performance table to check what flows are required for the selected therapy settings. |  |
| 13 | To deliver O₂% equal to 40% |  |
| 14 | and PEEP of 7,5cmH₂O |  |
| 15 | Adjust the PEEP valve. |  |
| 16 | Check line B is occluded using the blue pin. |  |
| 17 | Connect line A to the flow meter ensuring it is securely fitted. |  |
| 18 | Activate the flow according to the table. |  |
| 19 | Ventumask to be used only with oxygen flow meters capable of delivering the flows indicated in the table. |  |
| 20 | Fit the mask using the harness provided. |  |
| 21 | Check that the harness is comfortable for the patient. |  |
| 22 | Read just until desired fit. |  |
| 23 | You may fold over straps if they are obstructing. |  |
| 24 | Delivering O₂% greater than 40% Example of therapy settings O₂% = 50% PEEP = 7,5 cmH₂O |  |
| 25 | Additional oxygen flow required in line B |  |
| 26 | Remove the pin from line B. |  |
| 27 | Connect line B to the flow meter ensuring it is securely fitted. |  |
| 28 | Activate the flow according to the table. |  |
| 29 | Changing the PEEP value during therapy |  |
| 30 | Adjust the PEEP valve. |  |
| 31 | Adjust the flows according to the performance table, if required. |  |
| 32 | Delivering O₂% equal to 100% |  |
| 33 | Deactivate the Venturi driver by occluding the air inlet with the blue collar. |  |
| 34 | When the blue collar occludes the air inlet, the flow to the patient will be the total of lines A and B. |  |
| 35 | The minimum oxygen flow rate to the mask should be at least 42 litres per minute. |  |
| 36 | Removing the Ventumask |  |
| 38 | Undo the harness and remove the mask. |  |
| 39 | Turn off the flow. |  |
| 40 | Intersurgical Quality, innovation and choice Interact with us |  |