VHB20 humidifier script

|  | English |  |
| --- | --- | --- |
| 1 | Setting up a VHB20 Series Humidifier |  |
| 2 | You will need |  |
| 3 | Humidifier |  |
| 4 | Humidification chamber |  |
| 5 | Breathing system |  |
| 6 | Water (for autofill chambers only) |  |
| 7 | Sensor cable |  |
| 8 | Heated wire cable |  |
| 9 | Filters |  |
| 10 | Connecting the components |  |
| 11 | Place the humidifier into the humidifier bracket and connect to the mains. |  |
| 12 | Place the humidification chamber onto the humidifier and ensure that it clicks into place. |  |
| 13 | Ensure that the top of the humidification chamber is placed lower than the patient airway entry position. |  |
| 14 | Add water to the humidification chamber. In this case we are using an auto fill chamber. |  |
| 15 | To protect the ventilator we are using the Filta-guard™ on the expiratory limb. |  |
| 16 | And a high efficiency Air-Guard Clear filter on the inspiratory limb. |  |
| 17 | Remove the cover and attach the short inspiratory limb. |  |
| 18 | Connect the breathing system. The connector with the temperature port goes on the chamber. |  |
| 19 | The other end connects to the expiratory filter. |  |
| 20 | You may now like to stow the system while you connect the heated wire adaptor and temperature probes. |  |
| 21 | Connect the blue temperature probe adaptor and both probes. |  |
| 22 | Connect the red heated wire adaptor and both power plugs. |  |
| 23 | Connect the clover leaf to the inspiratory limb. |  |
| 24 | The oval connector goes to the expiratory limb if using a dual heated wire system. |  |
| 25 | Setting up the humidifier |  |
| 26 | Start and Select Mode |  |
| 27 | Power on the humidifier. |  |
| 28 | Select to use Previous or Default Settings. |  |
| 29 | Select Invasive Mode or Non-invasive Mode (In case of “Default Settings” in step 2) |  |
| 30 | To switch between the Chamber and Patient end temperature press the up arrow. |  |
| 31 | Patient end |  |
| 32 | Chamber outlet |  |
| 33 | Set Mode |  |
| 34 | While running, press and hold @ to switch between Invasive mode & Non-invasive modes. |  |
| 35 | Single Heated Wire Breathing System Set-up |  |
| 36 | Alarm occurs with red limb displayed on the screen. |  |
| 38 | Press @ to confirm, the limb of the icon will turn grey and the device will continue running. |  |
| 39 | This summarises the basic set-up. Please refer to the instructions for use for further information. |  |
| 40 | Quality, innovation and choice |  |
| 41 | Interact with us |  |
| 42 | www.intersurgical.com |  |
| 43 |  |  |