Oxygen Therapy Training Video

The various sub headings, such as 'Features' 'Method of Use' and 'Caution' should be in a standard format.

|  | English | Your Language |
| --- | --- | --- |
| 1 | Recent papers have shown that amongst practicing trainee doctors and nurses, there is a lack of knowledge about oxygen therapy devices and their correct use. There are also many false beliefs, with the result that many patients on general wards are incorrectly treated. |  |
| 2 | This video covers the main types of device commonly used for delivering oxygen therapy, which enable clinicians to comply with guidelines published worldwide such as the British Thoracic Society Guidelines and the American Association for Respiratory Care Guideline We can now look at how these devices work, when and how they are used, features of the devices, and precautions associated with their use. We hope this will increase the level of knowledge in this important discipline. |  |
| 3 | The devices for delivery of oxygen can be divided into two groups depending on whether they deliver a proportion, or all of the patient’s ventilatory requirement. |  |
| 4 | Variable oxygen concentration devices, also known as low flow devices, deliver oxygen at flows lower than the patient’s inspiratory requirements, so additional room air is also breathed in. |  |
| 5 | Fixed oxygen concentration devices, also known as high flow devices, deliver flows that meet the patient’s total inspiratory requirements. |  |
| 6 | Firstly let’s look at variable oxygen concentration devices and the various device options that are available. |  |
| 7 | Nasal Cannulae are one of the most common ways of delivering oxygen therapy. A Nasal Cannula consists of two soft prongs attached to the oxygen supply tubing. First attach the oxygen tubing to the oxygen flow meter and set the flow rate to the appropriate flow for the paitent. The prongs are inserted into the patient’s nares, the tubing is looped over the patient’s ears and the toggle is adjusted to ensure a good fit. Oxygen flows from the cannula into the patients nasopharynx, which acts as an anatomic reservoir. |  |
| 8 | Nasal cannulae are available with different prong shapes and are lightweight and generally comfortable. Oxygen can be administered at 6L/min or less and Flow rates of 4 L/min or less do not need to be humidified . Nasal cannulae can be used at very low flow rates with children, and mouth breathers can benefit from this device. Patients are able to eat, drink and speak whilst receiving oxygen therapy however the use of a nasal cannulae requires the upper airway to be free of obstruction |  |
| ## 9 | Nasal cannulae are contraindicated in patients with nasal obstruction, for example nasal polyps. |  |
| 10 | The concentration of oxygen received by the patient will vary as previously mentioned Oxygen concentration is affected by the selected oxygen flow rate, and by the patient’s tidal volume and breathing pattern. |  |
| 11 | There are various sizes available for adults, paediatrics and neonates and a variety of prong shapes are also available to suit varying circumstances. |  |
| ## 12 | Straight prongs are the original design and these are still the most frequently used in hospitals. |  |
| ##### 13 | Curved prongs are designed to direct the flow of oxygen posteriorly rather than upward towards the frontal sinuses. |  |
| ## 14 | ## Flared prongs have an increased diameter towards the tip. This has the effect of slowing down the gas as it enters the nose making it more comfortable for the patient. |  |
| ## 15 | ## Curved and flared prongs combine the best features of the two options, releasing oxygen more slowly and away from the frontal sinuses. |  |
| ## 16 | ## Nasal cannulae can be supplied with earguards which can benefit long term users of nasal cannulae by preventing sores on and around the ears |  |
| 17 | Nasal cannula are also available which allow sampling of exhaled carbon dioxide in non-intubated patients and the Intersurgical devices with this feature are called Sentri nasal cannula. These devices are suitable for patients undergoing conscious sedation where respiratory depression needs to be monitored. The Sentri nasal cannula is fitted to the patient in the same manner as a conventional nasal cannula. A monitoring line is then attached and connected to the capnograph. The Sentri nasal cannula allows delivery of oxygen through one nasal prong and sampling of exhaled gas through the other prong. The end tidal CO2 values are comparable to those achieved with intubated patients. |  |
| 18 | Now we can look at medium concentration oxygen masks which provide another option for the delivery of low flow oxygen therapy. These may be used for concentrations higher than those that can be achieved with nasal cannula or when the nares are unavailable. |  |
| ## 19 | ## These masks are available in adult and paediatric sizes. |  |
| 20 | The arrows demonstrate the manner in which gases are inhaled and exhaled by the patient. The red arrows represent the oxygen delivered to the mask. The orange arrows represent ambient air which is drawn in through the mask vents as the patients inhales and the blue arrows represent exhalation. The concentration of oxygen received by the patient will vary, however, as a guideline a flow rate of 5-8L/min will deliver concentrations between 35 and 50% oxygen to most adult patients. Oxygen concentration is affected by the selected oxygen flow rate, and by the patient’s tidal volume and breathing pattern. |  |
| 21 | One end of the oxygen tube is attached to the spigot on the mask and the other to the oxygen flowmeter. The flowmeter is set to give the prescribed oxygen concentration. The mask is then fitted over the patient’s nose and mouth and the elastic tightened, to give a secure comfortable fit. The elastic historically has been secured over the ears, however new generation Intersurgical EcoLite oxygen masks, can also be secured under the ears to avoid the occurrence of pressure sores. It is important to use a mask which obtains a good seal on the patients nose to prevent oxygen entering the eyes. A poorly fitting mask can expose the eye to a high flow of dry oxygen which can result in exposure keratopathy which then may lead to corneal perforation. |  |
| 22 | Caution must be taken when using these masks to ensure that the flow is at least 5L/min to wash carbon dioxide out of the mask and prevent rebreathing Aspiration of vomitus is also more likely when a mask is in place. |  |
| 23 | Medium concentration masks such as the Intersurgical Sentri mask, are also available which feature the ability to measure end-tidal CO2. This is beneficial for patients undergoing conscious sedation where respiratory depression should be monitored. A monitoring line can be attached to the mask with a luer lock connector and the other end is attached to a capnograph. |  |
| 24 | The final variable concentration oxygen mask we are looking at is the high concentration mask, also known as the non re-breathing mask and is usually used when high concentrations of oxygen are required, for example with the critically ill or unstable patient. |  |
| 25 | The patient breathes from the mask and the inspiratory valve opens so that oxygen is taken from the reservoir bag. The patient breathes out through the expiratory valve. |  |
| 26 | Ideally the patient does not breath any room air, however the fit of the mask and the patient’s breathing pattern will affect the amount of room air entrained and the oxygen concentration delivered. Limitations of the mask fit associated with older mask designs, meant that high concentration oxygen masks delivered up to 80% oxygen, with flows set between 10 & 15 L/min The improved fit achieved with the new generation Intersurgical EcoLite masks minimises the amount of air that is entrained through the sides of the mask and hence oxygen concentrations greater than 80% can be delivered to the patient. |  |
| 27 | This improved performance however makes it necessary to include mask safety valves. In the event of the oxygen supply failing and the reservoir bag emptying, the safety valves will open, allowing air to enter the mask and the patient is still able to breath. |  |
| 28 | When using the high concentration oxygen mask, First attach the oxygen tubing to the oxygen flowmeter and set the flow rate to 15 litres per minute. Occlude the inspiratory valve, being careful not to damage the valve and allow the reservoir bag to fill. Squeeze the reservoir bag to test the patency of the valve between the reservoir bag and the mask. If the reservoir will not empty discard the mask, choose another and test again. Refill the reservoir bag and place the mask on the patients face obtaining a tight but comfortable fit. . Adjust the oxygen flow rate so that the bag deflates by no more than one third, at the end of inspiration. |  |
| 29 | - When a mask is in place be aware that aspiration of vomitus is more likely . |  |
| 30 | An additional feature found in some high concentration oxygen masks is the Respi-Check breathing indicator. Respiratory rate should be recorded in anyone requiring oxygen therapy and yet it is often inaccurately estimated and poorly recorded. The Respi-check features a ball contained in a clear plastic tube which acts as a visual indicator of respiratory rate. The flow of oxygen into the mask pushes the ball to the top of the tube. The negative pressure generated by the patients inspiratory effort pulls the ball to the bottom of the tube, giving an instant visual notification of each breath taken. |  |
| 31 | Now let's look at the Fixed oxygen concentration devices, also known as High flow devices. These provide sufficient flow of gas to meet all of the patients minute ventilation requirements. |  |
| 32 | Firstly why would you use a fixed oxygen concentration device? These devices might be more appropriate for patients with chronic lung disease and CO2 retention, whose ventilation is dependent on hypoxic drive. This is an overemphasised condition where the usual stimulus for respiration, which is the maintenance of normal CO2 levels, is lost due to chronically high CO2 levels. Respiration is then driven by hypoxia. If high concentrations of Oxygen are given to these patients and their PaO2 increases, this hypoxic respiratory drive, can be lost. Hypoventilation may occur and the CO2 will become very high causing coma and possibly death. |  |
| 33 | A high flow device in which the oxygen concentration can be carefully controlled, and guided by blood gas measurements, is more appropriate for these patients. |  |
| 34 | The inspired oxygen concentration is determined by the design of the device and not by the patient’s rate and depth of breathing. |  |
| 35 | This is because venturi devices accurately deliver pre-determined oxygen concentrations at flows that are high enough to meet the inspiratory flow rates of most adult patients. The gas flow to the patient is increased by entraining air into the device using the venturi effect. Accurate oxygen concentrations are delivered by the mixing of entrained air and oxygen. |  |
| 36 | Masks are available with interchangeable, colour coded, venturi valves for different fixed oxygen concentrations, for example 35%. |  |
| 37 | When using a venturi mask, select a venturi valve of the prescribed oxygen concentration and attach it to the mask. Attach one end of the oxygen tubing to the valve spigot and the other end to the oxygen flowmeter. Set the flowmeter to the flow indicated on the valve. Place the mask over the patient’s nose and mouth to obtain a secure and comfortable fit. |  |
| 38 | It is important to be aware that at oxygen requirements above 40%, the valve may not generate enough total flow to meet high inspiratory demands. There is also a widespread misconception that the oxygen concentration varies with the oxygen flow rate. If the oxygen flow rate is increased, the total flow to the patient is increased but the oxygen concentration stays the same. For example the recommended flow rate for a 35% venturi valve is 8L/min. This delivers to the patient a flow rate of 45L/min and an oxygen concentration of 35%. Increasing the flow rate to 10L/min increases the total flow to the patient but the oxygen concentration delivered remains at 35%. |  |
| 39 | Fixed venturi valves can also be incorporated into Oxygen Recovery T-Piece Kits. The most popular is the 40% valve. These are used to provide spontaneously breathing, intubated patients with supplementary oxygen during transport and in the Recovery Room. |  |
| 40 | Attach one end of the oxygen tubing to the valve spigot and the other to the oxygen flowmeter and set the flowmeter to the flow indicated on the valve. Connect the T_Piece to the 15mm connector of the tracheal tube or supraglottic airway device. |  |
| ## 41 | ## A disadvantage of fixed concentration venturi valves is that in order to change the concentration delivered to the patient it is necessary to change the valve. ## This can be overcome by using an adjustable venturi valve. |  |
| 42 | With an adjustable venturi valve, the rotation of the sleeve increases or decreases the size of the air entrainment window; this in turn changes the oxygen concentration. This means that the same valve can deliver a wide range of concentrations. |  |
| 43 | As the valve does not have to be replaced each time the prescribed oxygen concentration changes, the changes can be made quickly. |  |
| 44 | It also means that different valves do not have to be held in stock by the hospital. |  |
| 45 | The Silente Adjustable Venturi, can be set to concentrations between 26 & 50% and incorporates a Twist/Lock mechanism to prevent inadvertent concentration changes. |  |
| 46 | As with a fixed valve attach one end of the oxygen tubing to the valve spigot and the other to the oxygen flow meter. To Select the desired oxygen concentration, ensure the range selector sleeve is in the appropriate position |  |
| 47 | If the desired concentration is between 26 and 35%, then ensure the oxygen range selector is set at 26-35 If the desired concentration is between 40 and 50%, rotate the oxygen concentration selector to 35% and then pull and rotate the oxygen range selector to select the 40-50 range. |  |
| 48 | Select the desired oxygen setting, by rotating the oxygen concentration selector. Set the oxygen flow meter to the flow indicated adjacent to the oxygen concentration. |  |
| 49 | The Intersurgical EcoLite range of masks represent a new development in oxygen therapy devices and address many of the limitations found in earlier designs. The materials used have a significantly lower impact on the environment. EcoLite oxygen masks have a lightweight design, incorporating a soft face seal, which provides a high level of patient comfort. Masks seal to the contours of different face shapes, to maximise delivery of oxygen to the patient and prevent oxygen from entering the eyes. EcoLite oxygen masks have also been designed to address tissue viability requirements, by reducing the risk of pressure sores developing. |  |
| ## 50 | The final section covers the different types of humidifier that deliver cool “humidification”. |  |
| 51 | Fink and Hunt in Clinical Practice in Respiratory Care describe the two most commonly used devices in this category as follows: |  |
| 52a | “A humidifier is a device that adds molecular water to gas, whereas a nebuliser produces an aerosol or suspension of particles in gas”. |  |
| 52b | Bubble Humidifiers use the bubble through humidification process to bring dry oxygen gas to ambient levels of humidity. The gas is directed into the water bottle where it is broken up into small bubbles, these gain humidity as they rise to the surface of the water. |  |
| 53 | The absolute humidity of the gas varies with flow rate, water temperature, bottle size and volume of water in the bottle. |  |
| 54 | Bubble humidifiers are most commonly used with low flow devices such as nasal cannulae and medium concentration oxygen masks. |  |
| 55 | Some models incorporate a pressure relief valve designed to prevent excessive pressures building up in the bottle should the oxygen tubing become occluded. Operation of the relief valve may also activate an audible warning drawing the attention of the user to this problem. |  |
| 56 | The head of the bubble humidifier may also incorporate a “universal” screw thread; this permits the use of a bottle, which can be refilled with sterile water, or the use of some pre-filled sterile water bottles. |  |
| 57 | The flowmeter should be attached to an appropriate oxygen supply. |  |
| 58 | Fill the bottle with sterile water using an aseptic technique and fit to the humidifier head. or fit a pre-filled sterile water bottle. |  |
| 59 | The humidifier should then be attached to the flowmeter. |  |
| 60 | To test the pressure relief valve, adjust the flowmeter to 4 L/min and occlude the outlet. Verify that the relief valve operates within 5 seconds. |  |
| 61 | Reset the oxygen flowmeter to the appropriate flow for the patient, |  |
| 62 | Attach the oxygen tubing to the bubble humidifier and the delivery device and fit to the patient. |  |
| 63 | The other method of delivering cold humidification is by using a humidifier nebuliser. |  |
| 64 | Humidifier Nebulisers use the venturi principle to create a low-pressure area at the end of a tube. The other end of the tube extends into a bottle containing sterile water. Water is drawn up the tube and is broken into an aerosol of small droplets by the oxygen stream. This same low-pressure area, also entrains air from the atmosphere, to dilute the oxygen to the desired level. The air/oxygen mixture, containing the water droplets, then passes to the patient. |  |
| 65 | Humidifier nebulisers can deliver between 24 and 50mgH20/L of gas, depending on the flow setting. Humidifier Nebulisers are most commonly used with devices such as aerosol, and tracheostomy masks. |  |
| 66 | It is important that these devices operate as quietly as possible to ensure a comfortable environment for the patient. They should also conform to international standards to ensure that the oxygen concentrations are delivered accurately. |  |
| 67 | Some devices, for example the AquaMist, are fitted with a tamper-proof control to prevent inadvertent changes to oxygen concentrations. The head of the AquaMist nebuliser humidifier incorporates a “universal” screw thread; this permits the use of a bottle, which can be refilled with sterile water, or the use of some pre-filled sterile water bottles. |  |
| 68 | The flowmeter should be attached to an appropriate oxygen supply. Fill the bottle with sterile water using an aseptic technique and fit to the humidifier head. The humidifier nebuliser should then be attached to the flowmeter. Attach 22mm Flextube to the outlet of the humidifier nebuliser It is important to position the tubing so that the lowest point is below the patient and to regularly drain any condensate, to prevent any condensed water from running into the patient or occluding the Flextube. Attach the aerosol mask and then adjust the oxygen concentration to the required level |  |
| 69 | Set the flowmeter to provide the required total flow to the patient. The Intersurgical AquaMist makes this simple by indicating the appropriate flow above each oxygen setting. Then fit the delivery device to the patient. An aerosol mask should be placed over the patient’s nose and mouth and the elastic is adjusted to give a secure and comfortable fit. When using a tracheostomy mask, the mask is positioned over the patients tracheostomy site and secured in place using an adjustable elastic strap around the patients neck. |  |
| 70 | There is no international standard for flowmeter connections and they can vary by country and manufacturer. The Aquaflow and and Aquamist both incorporate a DISS oxygen connection, which is the most common type. Flowmeter adaptors are available to fit other types of flowmeter. For example the M12 |  |
| 71 | These devices should not be used in ventilator circuits, as the air entrainment port, will introduce a leak into the system. These devices should also only be filled with sterile fluids. that are changed every day. Any residual fluid, should be drained from the bottle, and discarded before refilling. |  |
| 72 | To summarise, we have looked at many different oxygen therapy devices including :- Variable oxygen concentration devices Fixed oxygen concentration devices And Humidification devices and we hope this has helped to improve understanding in this important discipline. |  |
| 73 | Written Text on screen |  |
| 74 | Intersurgical (use logo) Presents Oxygen Therapy |  |
| 75 | While the information contained within this video was considered correct by Intersurgical at the time of production, it in no way qualifies anyone to undertake any medical activity. Intersurgical can accept no responsibility for the medical practices of persons who directly or indirectly reference this video. |  |
| 76 | Title: Oxygen Therapy Devices |  |
| 77 | BTS Guidelines for the emergency use of oxygen in adult patients Thorax 2008, Vol 63, Supplement VI. American Association for Respiratory Care Guideline: Oxygen Therapy for adults in the acute care facility, Respiratory Care, June 2002. |  |
| 78 | Covered in this video - How they work? - When and how they are used? - Device Features - Precautions |  |
| 79 | Graphic Oxygen Therapy Devices Variable Oxygen Concentration Fixed Oxygen Concentration Nasl Cannula Medium Concentration Masks High Concentration Masks Fixed Venturi Valves Adjustable Venturi Valves Oxygen Recovery T-piece Humidifier Nebulisers Humidification Devices Bubble Humidifier |  |
| 80 | Title: Variable Oxygen Concentration Devices |  |
| 81 | Title: Fixed Oxygen Concentration Devices |  |
| 82 | Title: Nasal Cannulae Method of Use |  |
| 83 | O2 flows from the cannulae into the patients nasopharynx. |  |
| 84 | Title: Nasal Cannulae Features - Available with different prong shapes. - Lightweight and generally comfortable - Oxygen can be administered at 6L/min or less. - Flow rates of 4 L/min or less do not need to be humidified - Can be used at very low flow rates with children - Mouth breathers can benefit from this device. - Patients are able to eat, drink and speak. - Needs upper airway to be free of obstruction. |  |
| 85 | Title: Nasal Cannulae Caution Nasal cannulae are contraindicated in patients with nasal obstruction, for example nasal polyps. |  |
| 86 | Oxygen concentration received by the patient will vary. Oxygen concentration is affected by the selected oxygen flow rate, and by the patient’s tidal volume and breathing pattern. Title: Nasal Cannulae Oxygen concentration received by the patient will vary. 1-2 L/min gives approximately 24-28% O2 2-4 L/min gives approximately 28-34% O2 4-6 L/min gives approximately 36-44% O2 |  |
| 87 | Title: Nasal Cannulae Size Shape |  |
| 88 | Title: Medium Concentration Oxygen Masks Features - Achieve higher O2 concentrations than nasal cannulae. - Suitable when nares are unavailable. |  |
| 89 | Title: Medium Concentration Oxygen Masks |  |
| 90 | Title: Medium Concentration Oxygen Masks Oxygen concentration is affected by the selected oxygen flow rate, and by the patient’s tidal volume and breathing pattern. Flow rate L/min O2 concentration 5 L/min will deliver approximately 35% O2 6 L/min will deliver approximately 40% O2 8 L/min will deliver approximately 50% O2 |  |
| 91 | Title: Medium Concentration Oxygen Masks Caution - Flow must be at least 5L/min - Aspiration of vomitus is more likely when a mask is in place. |  |
| 92 | Title: Medium Concentration Oxygen Masks |  |
| - 93 | - Title: High Concentration Oxygen Masks |  |
| ## 94 | ## Title: High Concentration Oxygen Masks A higher oxygen concentration is achievable with the EcoLite mask PVC mask EcoLite mask ## Features - Mask fit and breathing patters affect the O2 concentration Able to deliver oxygen concentrations greater - than 80% oxygen. |  |
| ## 95 | ## Title: High Concentration Oxygen Masks Features Safety valves, allowing air to enter the mask if the oxygen supply fails. |  |
| 96 | Title: High Concentration Oxygen Masks Method of Use |  |
| ## 97 | ## Title: High Concentration Oxygen Masks ## Caution - Aspiration of vomitus is more likely when a mask is in place. |  |
| 98 | Title: Fixed Oxygen Concentration Devices |  |
| 99 | Title: Fixed oxygen concentration venturi devices |  |
| 100 | Title: Adjustable oxygen concentration venturi valves |  |
| 101 | Title: Fixed Concentration Devices When? - Chronic Lung Disease patients whose ventilation is dependant on hypoxic drive - After prolonged high concentrations of O2, the hypoxic drive can be lost - Hypoventilation may occur, resulting in high CO2, risking coma and death. |  |
| 102 | Title: Fixed Concentration Devices |  |
| 103 | Title: Fixed Concentration Devices Inspired O2 concentration is determined by the design of the device and NOT the patient’s rate and depth of breathing. |  |
| 104 | Title: Fixed Concentration Devices Features High gas flow to the patient Venturi Effect |  |
| 105 | Title: Fixed Venturi Flow rates Oxygen flow meter - litres per minute (L/min) Oxygen Concentration Total flow rate to patient Litres per minute (L/min) Note: Figures in Bold indicate total flow rate to patient when using recommended oxygen flow. |  |
| 106 | Title: Fixed Venturi Recommended flow for a 35% venturi valve is 8L/min. Caution 8L/min O2 +37 L/min Air = Total flow rate 45 L/min. Caution Increasing the oxygen flow rate does NOT increase the oxygen concentration. 10 L/min O2 +46L/min Air = Total flow rate 56 L/min- |  |
| 107 | Title: Oxygen Recovery T-Piece Kits Features Provides spontaneously breathing, intubated patients with supplemetary oxygen during transport and in the recovery room. |  |
| 108 | Title: Oxygen Recovery T-Piece Kits Method of Use |  |
| 109 | Title: Adjustable Venturi Features Between 26% & 50% oxygen Twist & lock |  |
| 110 | Title: Adjustable Venturi Method of use |  |
| 111 | Title: Intersurgical EcoLite Oxygen Masks Environmental impact Using the simapro 67 life cycle assessment. Eco milli points score PVC Mask Eco Mask |  |
| 112 | Title: Humidification Devices |  |
| 113 | Title: Bubble Humidifier Flow rate To the patient |  |
| 114 | Title: Bubble Humidifier Humidity of the gas varies with: Flow rate Water temperature Bottle size Water Volume |  |
| 115 | Title: Bubble Humidifier Features - Commonly use with low flow devices - Pressure relief valve - Universal Screw head - Can be used with some prefilled sterile water bottles |  |
| 116 | Title: Bubble Humidifier Method of use |  |
| 117 | Title: Humidifier Nebuliser Oxygen Entrained air To patient Water |  |
| 118 | Title: Humidifier Nebuliser Features - Deliver 24-50mg H2O per litre of gas - Most commonly used with aerosol and tracheaostomy masks. |  |
| 119 | Title: Humidifier Nebuliser Features - Should Operate as quietly as possible - Should conform to international standards to ensure accurate O2 concentrations |  |
| 120 | Title: Humidifier Nebuliser Features - Some are tamperproof - Head may incorporate a Universal Screw Thread to take standard bottles |  |
| 121 | Title: Humidifier Nebuliser Method of Use Caution The lowest point of the flextube must be lower than the patient to allow condensation to drain safely. |  |
| 122 | Title: Humidifier Nebuliser |  |
| 123 | Title: Humidification Devices |  |
| 124 | Title: Humidification Devices Caution - Not for use in ventilator circuits - Sterile fluids only BOC M12 |  |
| 125 | Key Oxygen Exhalation Ambient air ETCO2 monitoring line |  |
| 126 | Covered Variable oxygen concentration devices Fixed concenrtation devices Humidification devices |  |

References

AUDIT UPDATE, British Thoracic Society emergency oxygen audits, B R O’Driscoll,1 L S Howard,2 C Bucknall,3 S A Welham,4 A G Davison,5 On behalf of the British Thoracic Society Thorax 2011;66:734e735

British Thoracic Society Guidelines for the emergency use of oxygen in adult patients. Thorax 2008, Volume 63, Supplement VI

American Association for Respiratory Care Guideline: Oxygen Therapy for adults in the acute care facility, Respiratory Care June 2002 Vol 47 no. 6.

Clinical Practice in Respiratory Care, James B. Fink, Gerald E. Hunt, Lippincott Williams & Wilkins, 1999