Oxygen Therapy Training Video

The various sub headings, such as 'Features' 'Method of Use' and 'Caution' should be in a standard format.

|  | English | Your Lanuage |
| --- | --- | --- |
| 1 | Recent papers have shown that amongst practicing trainee doctors and nurses, there is a lack of knowledge about oxygen therapy devices and their correct use. There are also many false beliefs, with the result that many patients on general wards are incorrectly treated. | 目前有文章显示，即便是一些训练有素的医护人员，对于氧气治疗设备及其正确的操作仍然缺乏一定的知识。 还有许多错误认知，得不到正确的治疗. |
| 2 | This video covers the main types of device commonly used for delivering oxygen therapy, which enable clinicians to comply with guidelines published worldwide such as the British Thoracic Society Guidelines and the American Association for Respiratory Care Guideline We can now look at how these devices work, when and how they are used, features of the devices, and precautions associated with their use. We hope this will increase the level of knowledge in this important discipline. | 本视频介绍了供氧治疗常用的主要设备类型，使临床医生能够遵守世界范围内发布的指南，如英国胸科协会指南和美国呼吸治疗协会指南。 现在，我们可以看下这些设备是如何工作的，何时以及如何使用它们，设备的特性 以及与使用相关的注意事项。 我们希望这可以提升这一重要学科的知识水平。 |
| 3 | The devices for delivery of oxygen can be divided into two groups depending on whether they deliver a proportion, or all of the patient’s ventilatory requirement. | 供氧设备可分为两组，一种是提供固定氧浓度，一种是根据患者的通气需求供氧。 |
| 4 | Variable oxygen concentration devices, also known as low flow devices, deliver oxygen at flows lower than the patient’s inspiratory requirements, so additional room air is also breathed in. | 可变氧浓度装置，也被称为低流量装置，输送的氧气流量低于病人吸气流量，因此额外的空气也会被吸入。 |
| 5 | Fixed oxygen concentration devices, also known as high flow devices, deliver flows that meet the patient’s total inspiratory requirements. | 固定氧浓度装置，也称为高流量装置，提供满足患者总吸气需求的流量。 |
| 6 | Firstly let’s look at variable oxygen concentration devices and the various device options that are available. | 首先让我们看看可变氧浓度装置和各种各样的可供选择的器械。 |
| 7 | Nasal Cannulae are one of the most common ways of delivering oxy gen therapy. A Nasal Cannula consists of two soft prongs attached to the oxygen supply tubing. First attach the oxygen tubing to the oxygen flow meter and set the flow rate to the appropriate flow for the paitent. The prongs are inserted into the patient’s nares, the tubing is looped over the patient’s ears and the toggle is adjusted to ensure a good fit. Oxygen flows from the cannula into the patients nasopharynx, which acts as an anatomic reservoir. | 鼻氧管是最常见的输氧方式之一。 鼻氧管由两个连接在氧气管上的软鼻塞组成。首先将氧气管接在氧气流量计上，并调整为合适流量。将鼻塞插= 入患者的鼻孔，然后将氧气管挂至耳后，调整拴扣确保适合患者。氧气通过鼻氧管流入患者鼻咽部，而鼻咽部刚好作为生理储氧库。 |
| 8 | Nasal cannulae are available with different prong shapes and are lightweight and generally comfortable. Oxygen can be administered at 6L/min or less and Flow rates of 4 L/min or less do not need to be humidified . Nasal cannulae can be used at very low flow rates with children, and mouth breathers can benefit from this device. Patients are able to eat, drink and speak whilst receiving oxygen therapy however the use of a nasal cannulae requires the upper airway to be free of obstruction | 鼻氧管有不同的鼻塞形状，重量轻，并且舒适。 可在6升/分钟或者更小流量下供氧 流量在4升/分钟或更小时不需要加湿 鼻氧管可以调到很小的流量用于儿童，同时适用于经口呼吸的患者。 患者在接受氧气治疗的同时不影响吃喝以及说话，然而鼻氧管在使用过程中要求上呼吸道不被阻塞。 |
| ## 9 | Nasal cannulae are contraindicated in patients with nasal obstruction, for example nasal polyps. | 鼻阻塞患者不能使用鼻氧管，例如鼻息肉 |
| 10 | The concentration of oxygen received by the patient will vary as previously mentioned. However, oxygen flows of 1-6L/min will deliver up to 44% of O2 to most adult patients. | 如前所述，病人接受的氧气浓度会有所不同。然而，在1-6升/分钟的氧气流量下可以为大多数成人患者提供高达44%的氧浓度）。 |
| 11 | There are various sizes available for adults, paediat-rics and neonates and a variety of prong shapes are also available to suit varying circumstances. | 成人、儿童和新生儿都有不同的尺寸，鼻塞也有各种各样的形状，以适应不同的情况。 |
| ## 12 | Straight prongs are the original design and these are still the most frequently used in hospitals. | 直鼻头是最初的设计，现在仍是医院里最常用的。 |
| ##### 13 | Curved prongs are designed to direct the flow of oxygen posteriorly rather than upward towards the frontal sinuses. | 弯鼻头设计是为了引导氧气向后方流动，而不是向上流向额窦。 |
| ## 14 | ## Flared prongs have an increased diameter towards the tip. This has the effect of slowing down the gas as it enters the nose making it more comfortable for the patient. | ## 喇叭鼻头前端直径增大，这有助于减缓气体进入鼻腔，使病人感觉更舒适。 |
| ## 15 | ## Curved and flared prongs combine the best features of the two options, releasing oxygen more slowly and away from the frontal sinuses. | 喇叭弯鼻头结合了这两者的最佳优势，既能降低氧气释放的速度又能引导氧气远离额窦 |
| ## 16 | ## Nasal cannulae can be supplied with earguards which can benefit long term users of nasal cannulae by preventing sores on and around the ears | 鼻氧管可配备耳套，对于长期使用鼻氧管的患者可预防耳朵以及耳周溃疡 |
| 17 | Nasal cannula are also available which allow sampling of exhaled carbon dioxide in non-intubated patients and the Intersurgical devices with this feature are called Sentri nasal cannula. These devices are suitable for patients undergoing conscious sedation where respiratory depression needs to be monitored. The Sentri nasal cannula is fitted to the patient in the same manner as a conventional nasal cannula. A monitoring line is then attached and connected to the capnograph. The Sentri nasal cannula allows delivery of oxygen through one nasal prong and sampling of exhaled gas through the other prong. The end tidal CO2 values are comparable to those achieved with intubated patients. | 鼻氧管也可用于非插管患者呼出的二氧化碳的取样，具有这种特征的（英特赛克）鼻氧管被称为Sentri鼻氧管。 这些设备适用于处于镇静状态并需要进行呼吸抑制监测的患者。 与传统的鼻腔插管一样，Sentri 鼻氧管的使用方法和传统的鼻氧管一致。将一条监测管连接到二氧化碳分析仪上。 Sentri鼻氧管通过一个鼻头输送氧气，通过另一个鼻头对呼出的气体进行取样。呼吸末二氧化碳的值与通过插管患者获取的值是一致的。 |
| 18 | Now we can look at medium concentration oxygen masks which pro vide another option for the delivery of low flow oxygen therapy. These may be used for concentrations higher than those that can be achieved with nasal cannula or when the nares are unavailable. | 现在我们可以看看中浓度的氧气面罩，它为低流量的氧气治疗提供了另一种选择。 这适用于对氧气浓度要求高于鼻氧管能够达到的最高浓度，或没有鼻氧管的情况下。 |
| ## 19 | ## These masks are available in adult and paediatric sizes. | 这些面罩有成人和儿童尺寸。 |
| 20 | The arrows demonstrate the manner in which gases are inhaled and exhaled by the patient. The red arrows represent the oxygen delivered to the mask. The orange arrows represent ambient air which is drawn in through the mask vents as the patients inhales and the blue arrows represent exhalation. The concentration of oxygen received by the patient will vary, however, as a guideline a flow rate of 5-8L/min will deliver concentrations between 35 and 50% oxygen to most adult patients. | 箭头表示病人吸入和呼出气体的方式。红色箭头表示输送给面罩的氧气。橙色箭头表示患者吸入时通过面罩通风口吸入的空气，蓝色箭头表示呼气。 患者接受的氧气浓度会有所不同，然而，在5-8升/分钟的流量下可以为大多数成人患者提供35%到50%的氧气浓度，这可以作为一个指南 |
| 21 | One end of the oxygen tube is attached to the spigot on the mask and the other to the oxygen flowmeter. The flowmeter is set to give the prescribed oxygen concentration. The mask is then fitted over the patient’s nose and mouth and the elastic tightened, to give a secure comfortable fit. The elastic historically has been secured over the ears, however new generation Intersurgical EcoLite oxygen masks, can also be secured under the ears to avoid the occurrence of pressure sores. It is important to use a mask which obtains a good seal on the patients nose to prevent oxygen entering the eyes. A poorly fitting mask can expose the eye to a high flow of dry oxygen which can result in exposure keratopathy which then may lead to corneal perforation. | 氧气管的一端与面罩上的接头相连，另一端与氧气流量计相连。 设定流量计，使其能够提供指定的氧浓度。 将面罩套在病人的鼻子和嘴巴上，然后收紧松紧带，保证安全和舒适。传统的松紧带往往固定在耳朵上，而新一代的英特赛克Ecolite氧气面罩，也可以固定在耳朵下，以避免压疮的出现。 选用一个密封性好的面罩非常重要，可以防止氧气进入眼睛。不合适的面罩会使眼睛暴露在高流量的干氧下，这可能导致暴露性角膜病变，最后可能导致角膜穿孔。 |
| 22 | Caution must be taken when using these masks to ensure that the flow is at least 5L/min to wash carbon dioxide out of the mask and prevent rebreathing Aspiration of vomitus is also more likely when a mask is in place. | 使用这些面罩时必须注意，确保流量至少为5L/min，这样才能将二氧化碳从口罩中排出去，并防止重复吸入。 戴面罩时也更有可能导致呕吐物的误吸。 |
| 23 | Medium concentration masks such as the Intersurgical Sentri mask, are also available which feature the ability to measure end-tidal CO2. This is beneficial for patients undergoing conscious sedation where respiratory depression should be monitored. A monitoring line can be attached to the mask with a luer lock connector and the other end is attached to a capnograph. | 能够监测呼吸末二氧化碳的中浓度氧气面罩也是有的，如英特赛克的Sentri面罩。这对于处于镇静状态并需要进行呼吸抑制监测的患者是十分有益的。监测管一端可以通过鲁尔接头连接到面罩上，另一端连接到二氧化碳分析仪。 |
| 24 | The final variable concentration oxygen mask we are looking at is the high concentration mask, also known as the non re-breathing mask and is usually used when high concentrations of oxygen are required, for example with the critically ill or unstable patient. | 最后一款可变浓度氧气面罩是高浓度氧气面罩，也被称为非重复呼吸面罩，通常在需要高浓度氧气时使用，例如危重病人或不稳定病人。 |
| 25 | The patient breathes from the mask and the inspiratory valve opens so that oxygen is taken from the reservoir bag. The patient breathes out through the expiratory valve. | 患者通过面罩呼吸，吸气阀打开，使氧气从储氧袋中吸出。 病人通过呼气阀呼气。 |
| 26 | Ideally the patient does not breath any room air, however the fit of the mask and the patient’s breathing pattern will affect the amount of room air entrained and the oxygen concentration delivered. Limitations of the mask fit associated with older mask designs, meant that high concentration oxygen masks delivered up to 80% oxygen, with flows set between 10 & 15 L/min The improved fit achieved with the new generation Intersurgical EcoLite masks minimises the amount of air that is entrained through the sides of the mask and hence concentrations approaching 90% oxygen can be delivered to the patient where flows are set between 10 & 15 L/min. . | 理想情况下，患者不会吸入任何室内空气，然而面罩的合适程度和患者的呼吸模式将会影响室内空气的夹带量和氧气的输送浓度。由于老款面罩的设计造成面罩贴合度的限制，意味着高浓度氧气面罩在10-15升/分钟的流量下氧气浓度最高可达80%。 英特赛克新一代的Ecolite面罩通过改善设计，将通过面罩两侧的空气量减到最小，因此在10-15升/分钟的流量下，到达患者的氧浓度可接近90%。 |
| 27 | This improved performance however makes it necessary to include mask safety valves. In the event of the oxygen supply failing a-nd the reservoir bag emptying, the safety valves will open, allowing air to enter the mask and the patient is still able to breath. | 这种改善使我们必须增加面罩安全阀门。当氧气供应不足，储氧袋排空时，安全阀将打开，允许空气进入面罩，患者仍能呼吸。 |
| 28 | When using the high concentration oxygen mask, First attach the oxygen tubing to the oxygen flowmeter and set the flow rate to 15 litres per minute. Occlude the inspiratory valve, being careful not to damage the valve and allow the reservoir bag to fill. Squeeze the reservoir bag to test the patency of the valve between the reservoir bag and the mask. If the reservoir will not empty discard the mask, choose another and test again. Refill the reservoir bag and place the mask on the patients face obtaining a tight but comfortable fit. . Adjust the oxygen flow rate so that the bag deflates by no more than one third, at the end of inspiration. | 使用高浓度氧气面罩时， 首先将氧气管接到氧气流量计上，并将流速设置为15升每分钟。 堵住吸入阀，注意不要损坏阀门，让储气袋充满。 挤压储氧袋来测试阀门在储氧袋和面罩之间是否通畅）。如果储氧袋没有排空则丢弃此面罩重新选择一个，然后再次进行测试。 将储氧袋重新充满，并将面罩戴在病人脸上，使其贴合并且舒适。 调整氧气流量，使储氧袋在吸气结束时放气不超过三分之一。 |
| 29 | - When a mask is in place be aware that aspiration of vomitus is more likely . | 需要注意戴面罩时更有可能导致呕吐物的误吸 |
| 30 | An additional feature found in some high concentration oxygen masks is the Respi-Check breathing indicator. Respiratory rate should be recorded in anyone requiring oxygen therapy and yet it is often inaccurately estimated and poorly recorded. The Respi-check features a ball contained in a clear plastic tube which - tube, giving an instant visual notification of each breath taken. | 一些高浓度氧气面罩的另一个特点是呼吸检查指示器。任何需要氧气治疗的人的呼吸频率都应该记录，但它往往是估计不准确的并且记录不良。呼吸检查指示器的特点是在一个透明的塑料管中有一个球，它可以作为呼吸频率的视觉指标。进入面罩的氧气流将球推到管子的顶部。病人 吸气时产生的负压将球拉到管子的底部，使每次呼吸都能立即看到。 |
| 31 | Now let's look at the Fixed oxygen concentration devices, also known as High flow devices. These provide sufficient flow of gas to meet all of the patients minute ventilation requirements. | 现在我们来看看固定氧浓度装置，也被称为高流量装置。 这些提供了足够的气体流量，以满足所有病人的通气需求。 |
| 32 | Firstly why would you use a fixed oxygen concentration device? These devices might be more appropriate for patients with chronic lung disease and CO2 retention, whose ventilation is dependent on hypoxic drive. This is an overemphasised condition where the usual stimulus for respiration, which is the maintenance of normal CO2 levels, is lost due to chronically high CO2 levels. Respiration is then driven by hypoxia. - If high concentrations of Oxygen are given to these patients and their PaO2 increases, this hypoxic respiratory drive, can be lost. Hypoventilation may occur and the CO2 will become very high causing coma and possibly death. | 首先，为什么要使用固定的氧浓度设备? 这些设备可能更适用于慢性肺病和二氧化碳滞留患者，他们的通气依赖于缺氧驱动。 这是一种过分强调的情况，在这种情况下，通常的呼吸刺激，即维持正常的二氧化碳水平，会由于长期的高二氧化碳水平而失去。 呼吸依靠缺氧驱动进行。 如果给这些病人高浓度的氧气，他们的血氧分压增加，这种低氧呼吸驱动力就会丧失。 通气不足可能发生，二氧化碳水平会变得非常高，导致昏迷甚至死亡。 |
| 33 | A high flow device in which the oxygen concentration can be carefully controlled, and guided by blood gas measurements, is more appropriate for these patients. | 一个高流量的设备由于它的血液气体测量可以精确控制和指导氧气浓度，所以更适合这些病人。 |
| 34 | The inspired oxygen concentration is determined by the design of the device and not by the patient’s rate and depth of breathing. | 吸入氧浓度是由装置的设计决定的，而不是由病人的呼吸频率和深度决定的。 |
| 35 | This is because venturi devices accurately deliver pre-determined oxygen concentrations at flows that are high enough to meet the inspiratory flow rates of most adult patients. The gas flow to the patient is increased by entraining air into the device using the venturi effect. Accurate oxygen concentrations are delivered by the mixing of entrained air and oxygen. | 这是因为文丘里装置能够准确地在足够高的流量下提供预先确定的氧气浓度，以满足大多数成年患者的吸气流速。 通过使用文丘里效应将空气吸入装置，增加了流向病人的气体流量。夹带的空气和氧气混合气体提供精确的氧气浓度。 |
| 36 | Masks are available with interchangeable, colour coded, venturi valves for different fixed oxygen concentrations, for example 35%. | 面罩可配不同的文氏阀，不同颜色标识的文氏阀代表不同的固定氧浓度例如35% |
| 37 | When using a venturi mask, select a venturi valve of the prescribed oxygen concentration and attach it to the mask. Attach one end of the oxygen tubing to the valve spigot and the other end to the oxygen flowmeter. Set the flowmeter to the flow indicated on the valve. Place the mask over the patient’s nose and mouth to obtain a secure and comfortable fit. | 当使用文丘里面罩时，选择固定氧浓度的文氏阀并将其连接到面罩上。 将氧气管的一端连接到阀门插管上，另一端连接到氧气流量计上。 将流量计设置为阀门上指示的流量。 将面罩盖住病人的鼻子和嘴巴，以确保安全舒适。 |
| 38 | It is important to be aware that at oxygen requirements above 40%, the valve may not generate enough total flow to meet high inspiratory demands. There is also a widespread misconception that the oxygen concentration varies with the oxygen flow rate. If the oxygen flow rate is increased, the total flow to the patient is increased but the oxygen concentration stays the same. For example the recommended flow rate for a 35% venturi valve is 8L/min. This delivers to the patient a flow rate of 45L/min and an oxygen concentration of 35%. Increasing the flow rate to 10L/min increases the total flow to the patient but the oxygen concentration delivered remains at 35%. | 需要注意的是，在氧浓度需求超过40%的情况下，阀门可能不会产生足够的流量来满足较高的吸气需求。 还有一种普遍的误解认为氧浓度随氧流量的变化而变化。如果氧流量增加，流向患者的总流量会增加，但氧浓度保持不变。 例如，35%文丘里阀门的推荐流量为8L/min。这为患者提供了45L/min的流量和35%的氧气浓度。 将流量增加到10L/min，可以增加流向患者的总流量，但输送的氧气浓度依然保持在35%。 |
| 39 | Fixed venturi valves can also be incorporated into Oxygen Recovery T-Piece Kits. The most popular is the 40% valve. These are used to provide spontaneously breathing, intubated patients with supplementary oxygen during transport and in the Recovery Room. | 固定文氏阀也可以在氧气康复T型组件中组合使用，使用最广泛的是40%浓度文氏阀。 这是用来为自主呼吸的插管患者在转移过程中或恢复室提供氧气。 |
| 40 | Attach one end of the oxygen tubing to the valve spigot and the other to the oxygen flowmeter and set the flowmeter to the flow indicated on the valve. Connect the T_Piece to the 15mm connector of the tracheal tube or supraglottic airway device. | 将氧气管的一端连接到阀门插管上，另一端连接到氧气流量计上 将流量计设置为阀门上指示的流量。 将T型接口连接到气管导管或声门上气道装置的15mm接头上。 |
| ## 41 | ## A disadvantage of fixed concentration venturi valves is that in order to change the concentration delivered to the patient it is necessary to change the valve. ## This can be overcome by using an adjustable venturi valve. | 固定浓度文氏阀的一个缺点是，如果要改变到达患者的氧浓度必须更换阀门。 这可以通过使用可调文氏阀门来解决。 |
| 42 | With an adjustable venturi valve, the rotation of the sleeve increases or decreases the size of the air entrainment window; this in turn changes the oxygen concentration. This means that the same valve can deliver a wide range of concentrations. | 通过可调文氏阀套筒的旋转可以增加或减少空气输送口的大小，从而改变氧气浓度。这意味着同一个阀门可以提供不同的氧气浓度。 |
| 43 | As the valve does not have to be replaced each time the prescribed oxygen concentration changes, the changes can be made quickly. | 由于不需要在每次指定的氧浓度发生变化时更换阀门，因此可以快速及时地改变氧浓度。 |
| 44 | It also means that different valves do not have to be held in stock by the hospital. | 这意味着医院不需要备很多不同阀门的库存。 |
| 45 | The Silente Adjustable Venturi, can be set to concentrations between 26 & 50% and incorporates a Twist/Lock mechanism to prevent inadvertent concentration changes. | Silente可调文氏阀，可以将浓度设置在26%到50%之间，并有一个扭转/锁机制，以防止无意中的浓度变化。 |
| 46 | As with a fixed valve attach one end of the oxygen tubing to the valve spigot and the other to the oxygen flow meter. To Select the desired oxygen concentration, ensure the range selector sleeve is in the appropriate position | 与固定阀门一样，将氧气管的一端连接到阀门插口上，另一端连接到氧气流量计上。 选择所需的氧浓度，请确保范围选择器处于正确的位置 |
| 47 | If the desired concentration is between 26 and 35%, then ensure the oxygen range selector is set at 26-35 If the desired concentration is between 40 and 50%, rotate the oxygen concentration selector to 35% and then pull and rotate the oxygen range selector to select the 40-50 range. | 如果所需浓度在26- 35%之间，则确保氧气范围选择器设置为26-35 如果所需浓度在40- 50%之间，则确保氧气范围选择器设置为40-50， |
| 48 | Select the desired oxygen setting, by rotating the oxygen concentration selector. Set the oxygen flow meter to the flow indicated adjacent to the oxygen concentration. | 通过旋转氧浓度选择器选择所需的氧浓度。将氧气流量计设置为与氧气浓度相邻的流量。 |
| 49 | The Intersurgical EcoLite range of masks represent a new development in oxygen therapy devices and address many of the limitations found in earlier designs. The materials used have a significantly lower impact on the environment. EcoLite oxygen masks have a lightweight design, incorporating a soft face seal, which provides a high level of patient comfort. Masks seal to the contours of different face shapes, to maximise delivery of oxygen to the patient and prevent oxygen from entering the eyes. EcoLite oxygen masks have also been designed to address tissue viability requirements, by reducing the risk of pressure sores developing. | 英特赛克Ecolite系列面罩代表了氧疗设备的新发展，并解决了许多早期设计的局限性。 所用的材料对环境的影响大大降低。 Ecolite氧气面罩轻量化设计，采用柔软的脸部密封，让患者更舒适。 面罩可以贴合不同脸型，最大限度地向病人输送氧气，并防止氧气进入眼睛。 Ecolite氧气面罩的设计通过减少压力疮的发生来解决组织活力的需求。 |
| ## 50 | The final section covers the different types of humidifier that deliver cool “humidification”. | 最后一节涵盖了不同类型的加湿器，它们提供了凉爽的“湿度”。 |
| 51 | Fink and Hunt in Clinical Practice in Respiratory Care describe the two most commonly used devices in this category as follows: | 在呼吸治疗的临床实践中芬克和亨特描述了这类设备中最常用的两种设备如下: |
| 52 | “A humidifier is a device that adds molecular water to gas, whereas a nebuliser produces an aerosol or suspension of particles in gas”. | 加湿器是一种向气体中添加水分子的装置，而雾化器则是在气体中制造喷雾或悬浮粒子的装置。 |
| 52 | Bubble Humidifiers use the bubble through humidification process to bring dry oxygen gas to ambient levels of humidity. The gas is directed into the water bottle where it is broken up into small bubbles, these gain humidity as they rise to the surface of the water. | 气泡加湿器在增湿过程中通过气泡将干燥的氧气增湿至室内空气的湿度水平。 这些气体直接进入水瓶，在那里被分解成小气泡，当它们上升到水面时，湿度就会增加。 |
| 53 | The absolute humidity of the gas varies with flow rate, water temperature, bottle size and volume of water in the bottle. | 气体的绝对湿度随以下情况而变化 流量 水温 瓶子尺寸 瓶中水量 |
| 54 | Bubble humidifiers are most commonly used with low flow devices such as nasal cannulae and medium concentration oxygen masks. | 气泡加湿器最常用于低流量器械如鼻氧管和中等浓度的氧气面罩。 |
| 55 | Some models incorporate a pressure relief valve designed to prevent excessive pressures building up in the bottle should the oxygen tubing become occluded. Operation of the relief valve may also activate an audible warning drawing the attention of the user to this problem. | 有些型号装有减压安全阀，用于防止氧气管堵塞时瓶内压力过大。 安全阀还可以触发警报，以引起用户的注意。 |
| 56 | The head of the bubble humidifier may also incorporate a “universal” screw thread; this permits the use of a bottle, which can be refilled with sterile water, or the use of some pre-filled sterile water bottles. | 气泡加湿器的头部也可装有“通用”螺纹;使其可以使用循环装满无菌水的瓶子，或者使用一些预先灌装的无菌水瓶。 |
| 57 | The flowmeter should be attached to an appropriate oxygen supply. | 流量计应与适当的供氧装置相连。 |
| 58 | Fill the bottle with sterile water using an aseptic technique and fit to the humidifier head or fit a pre-filled sterile water bottle. | 用无菌技术将无菌水装入瓶中并装好加湿器头，或直接装上预充无菌水瓶。 |
| 59 | The humidifier should then be attached to the flowmeter. | 加湿器应该连接到流量计上。 |
| 60 | To test the pressure relief valve, adjust the flowmeter to 4 L/min and occlude the outlet. Verify that the relief valve operates within 5 seconds. | 测试减压阀，将流量计调至4升/分钟，并堵塞出口。确认安全阀在5秒内启用。 |
| 61 | Reset the oxygen flowmeter to the appropriate flow for the patient, | 将氧气流量计重置为适合患者的流量， |
| 62 | Attach the oxygen tubing to the bubble humidifier and the delivery device and fit to the patient. | 将氧气管连接到气泡加湿器和输送装置上，适合患者使用。 |
| 63 | The other method of delivering cold humidification is by using a humidifier nebuliser. | 另一种增湿的方法是使用加湿喷雾器。 |
| 64 | Humidifier Nebulisers use the venturi principle to create a low-pressure area at the end of a tube. The other end of the tube extends into a bottle containing sterile water. Water is drawn up the tube and is broken into an aerosol of small droplets by the oxygen stream. This same low-pressure area, also entrains air from the atmosphere, to dilute the oxygen to the desired level. The air/oxygen mixture, containing the water droplets, then passes to the patient. | 加湿型喷雾器使用文丘里原理在管子末端创造一个低压区域。 管子的另一端伸入一个装有无菌水的瓶子。 水被吸入管中，被氧气流分解成气雾。 同样的低压区也会从大气中吸收空气，将氧气稀释到理想的浓度。 空气/氧气混合气体包含水滴，然后传递给病人 |
| 65 | Humidifier nebulisers can deliver between 24 and 50mgH20/L of gas, depending on the flow setting. Humidifier Nebulisers are most commonly used with devices such as aerosol, and tracheostomy masks. | 加湿型喷雾器每升气体可以提供24到50mg的水分，主要取决于流量设置。 加湿型喷雾器是最常用于雾化面罩和气切面具。 |
| 66 | It is important that these devices operate as quietly as possible to ensure a comfortable environment for the patient. They should also conform to international standards to ensure that the oxygen concentrations are delivered accurately. | 很重要的一点是这些设备操作非常安静可以确保病人有一个舒适的环境。 它们还应符合国际标准，以确保氧气浓度准确输送。 |
| 67 | Some devices, for example the AquaMist, are fitted with a tamper-proof control to prevent inadvertent changes to oxygen concentrations. As with bubble humidifiers, the head of the AquaMist nebuliser humidifier incorporates a “universal” screw thread; this permits the use of a bottle, which can be refilled with sterile water, or the use of some pre-filled sterile water bottles. | 有些设备，例如AquaMist，安装了防篡改的控制装置，以防止无意中改变氧浓度。 与气泡加湿器一样，AquaMist喷雾加湿器的头也采用了“通用”螺纹; 使其可以使用循环装满无菌水的瓶子，或者使用一些预先灌装的无菌水瓶。 |
| 68 | The flowmeter should be attached to an appropriate oxygen supply. Fill the bottle with sterile water using an aseptic technique and fit to the humidifier head. The humidifier nebuliser should then be attached to the flowmeter. Attach 22mm Flextube to the outlet of the humidifier nebuliser It is important to position the tubing so that the lowest point is below the patient and to regularly drain any condensate, to prevent any condensed water from running into the patient or occluding the Flextube. Attach the aerosol mask and then adjust the oxygen concentration to the required level | 流量计应与适当的氧气供应装置相连。 用无菌技术将瓶子装满无菌水，并与加湿器的头部相连。 喷雾加湿器应该连接到流量计。 将22mm可伸缩管连接到喷雾加湿器的出口 导管放置的位置很重要，其最低点要低于病人，并定期排出冷凝水，以防止冷凝水流入病人或堵塞可伸缩管。 安装雾化面罩，然后将氧气浓度调整到所需的水平 |
| 69 | Set the flowmeter to provide the required total flow to the patient. The Intersurgical AquaMist makes this simple by indicating the appropriate flow above each oxygen setting. Then fit the delivery device to the patient. An aerosol mask should be placed over the patient’s nose and mouth and the elastic is adjusted to give a secure and comfortable fit. When using a tracheostomy mask, the mask is positioned over the patients tracheostomy site and secured in place using an adjustable elastic strap around the patients neck. | 设置流量计，为患者提供所需的总流量。 英特赛克AquaMist通过指示每个氧气设置范围使之变得简单。 然后将装置连在病人端。 雾化面罩应放置在病人的鼻子和嘴巴上方，并调整松紧带保证安全和舒适。当使用气切面罩时，面罩放置在患者气切口处，用可调节的松紧带固定在患者颈部。 |
| 70 | There is no international standard for flowmeter connections and they can vary by country and manufacturer. The Aquaflow and and Aquamist both incorporate a DISS oxygen connection, which is the most common type. Flowmeter adaptors are available to fit other types of flowmeter. For example the M12 | 流量计的连接没有国际标准，根据国家和制造商的不同而变化。 Aquaflow和Aquamist都有DISS氧气接口，这是最常见的类型。 流量计适配器可适用于其他类型的流量计。例如M12 |
| 71 | These devices should not be used in ventilator circuits, as the air entrainment port, will introduce a leak into the system. These devices should also only be filled with sterile fluids that are changed every day. Any residual fluid, should be drained from the bottle, and discarded before refilling. | 这些装置不能用在呼吸回路中，因为吸气口会导致系统泄漏. 这些设备只能装无菌液体并且需要每天更换。在重新灌装前任何残留液体都应从瓶中抽出并丢弃。 |
| 72 | To summarise, we have looked at many different oxygen therapy devices including :- Variable oxygen concentration devices Fixed oxygen concentration devices And Humidification devices and we hope this has helped to improve understanding in this important discipline. | 综上所述，我们展示了许多不同的氧气治疗设备，包括: 可变氧浓度装置 固定氧浓度装置 增湿设备 我们希望这有助于增进对这一重要学科的理解。 |
| 73 | Written Text on screen |  |
| 74 | Intersurgical (use logo) Presents Oxygen Therapy | 英特赛克展示氧气治疗 |
| 75 | While the information contained within this video was considered correct by Intersurgical at the time of production, it in no way qualifies anyone to undertake any medical activity. Intersurgical can accept no responsibility for the medical practices of persons who directly or indirectly reference this programme. | 尽管在制作视频时，英特赛克认为视频中包含的信息是正确的，但这并不意味着任何人有资格从事任何医疗活动。 对于直接或间接引用本视频医疗行为的人员，英特赛克不承担任何责任。 |
| 76 | Title: Oxygen Therapy Devices | 氧气治疗设备 |
| 77 | BTS Guidelines for the emergency use of oxygen in adult patients Thorax 2008, Vol 63, Supplement VI. American Association for Respiratory Care Guideline: Oxygen Therapy for adults in the acute care facility, Respiratory Care, June 2002. | 《2008年成人胸腔病人紧急供氧BTS指南》 第63卷，补编VI 美国呼吸治疗协会指南:成人在急救护理中的氧气治疗，呼吸治疗，2002年6月。 |
| 78 | Covered in this video - How they work? - When and how they are used? - Device Features - Precautions | 如何工作的？ 何时使用以及如何使用？ 设备特征 注意事项 |
| 79 | Graphic Oxygen Therapy Devices Variable Oxygen Concentration Fixed Oxygen Concentration Nasl Cannula Medium Concentration Masks High Concentration Masks Fixed Venturi Valves Adjustable Venturi Valves Oxygen Recovery T-piece Humidifier Nebulisers Humidification Devices Bubble Humidifier | 图解 氧气治疗设备 可调氧气浓度 固定氧气浓度 鼻氧管 中浓度面罩 高浓度面罩 固定文氏阀 可调文氏阀 氧气康复T-型接口 喷雾型加湿器 加湿设备 气泡加湿器 |
| 80 | Title: Variable Oxygen Concentration Devices | 可调氧气浓度设备 |
| 81 | Title: Fixed Oxygen Concentration Devices | 固定氧气浓度设备 |
| 82 | Title: Nasal Cannulae Method of Use | 鼻氧管 使用方法 |
| 83 | O2 flows from the cannulae into the patients nasopharynx. | 氧气从鼻氧管通往病人鼻腔 |
| 84 | Title: Nasal Cannulae Features - Available with different prong shapes. - Lightweight and generally comfortable - Oxygen can be administered at 6L/min or less. - Flow rates of 4 L/min or less do not need to be humidified - Can be used at very low flow rates with children - Mouth breathers can benefit from this device. - Patients are able to eat, drink and speak. - Needs upper airway to be free of obstruction. | 鼻氧管 特征 有不同的鼻塞形状 重量轻且舒适 氧气流量可调至6L/分钟或者更小 流量小于等于4L/分钟无需加湿 儿童使用时可以调到很小流量 有利于经口呼吸的患者 病人可以吃，喝和说话 要求上呼吸道不被阻塞 |
| 85 | Title: Nasal Cannulae Caution Nasal cannulae are contraindicated in patients with nasal obstruction, for example nasal polyps. | 鼻氧管 注意 鼻阻塞患者不能使用鼻氧管，例如鼻息肉 |
| 86 | Title: Nasal Cannulae Oxygen concentration received by the patient will vary. 1-2 L/min gives approximately 24-28% O2 2-4 L/min gives approximately 28-34% O2 4-6 L/min gives approximately 36-44% O2 | 鼻氧管 患者接收到的氧浓度会变化： 1-2L/分钟提供氧浓度大约24-28% 2-4L/分钟提供氧浓度大约28-34% 4-6L/分钟提供氧浓度大约36-44% |
| 87 | Title: Nasal Cannulae Size Shape | 鼻氧管 大小 形状 |
| 88 | Title: Medium Concentration Oxygen Masks Features - Achieve higher O2 concentrations than nasal cannulae. - Suitable when nares are unavailable. | 中浓度氧气面罩 特征 到达的氧浓度比鼻氧管更高 适用于没有鼻氧管的情况下 |
| 89 | Title: Medium Concentration Oxygen Masks | 中浓度氧气面罩 |
| 90 | Title: Medium Concentration Oxygen Masks Flow rate L/min O2 concentration 5 L/min will deliver approximately 35% O2 6 L/min will deliver approximately 40% O2 8 L/min will deliver approximately 50% O2 | 中浓度氧气面罩 流量 L/分钟 氧气浓度 5 L/分钟提供氧浓度大约35% 6 L/分钟提供氧浓度大约40% 8 L/分钟提供氧浓度大约50% |
| 91 | Title: Medium Concentration Oxygen Masks Caution - Flow must be at least 5L/min - Aspiration of vomitus is more likely when a mask is in place. | 中浓度氧气面罩 注意 流量至少5L/分钟 戴面罩时也更有可能导致呕吐物的误吸 |
| 92 | Title: Medium Concentration Oxygen Masks | 中浓度氧气面罩 |
| - 93 | - Title: High Concentration Oxygen Masks | - 高浓度氧气面罩 |
| ## 94 | ## Title: High Concentration Oxygen Masks A higher oxygen concentration is achievable with the EcoLite mask PVC mask EcoLite mask ## Features - Mask fit and breathing patters affect the O2 concentration - Deliver close to 90% O2 | - 高浓度氧气面罩 Eoclite面罩能获得更高氧浓度 PVC面罩 Ecolite面罩 特征 面罩贴合度和呼吸模式影响氧气浓度 氧浓度接近90% |
| ## 95 | ## Title: High Concentration Oxygen Masks Features Safety valves, allowing air to enter the mask if the oxygen supply fails. | - 高浓度氧气面罩 特征 安全阀门。当氧气供应不足，允许空气进入面罩 |
| 96 | Title: High Concentration Oxygen Masks Method of Use | - 高浓度氧气面罩 |
| ## 97 | ## Title: High Concentration Oxygen Masks ## Caution - Aspiration of vomitus is more likely when a mask is in place. | - 高浓度氧气面罩 注意 戴面罩时也更有可能导致呕吐物的误吸 |
| 98 | Title: Fixed Oxygen Concentration Devices | 固定氧浓度设备 |
| 99 | Title: Fixed Concentration Devices When? - Chronic Lung Disease patients whose ventilation is dependant on hypoxic drive - After prolonged high concentrations of O2, the hypoxic drive can be lost - Hypoventilation may occur, resulting in high CO2, risking coma and death. | 固定氧浓度设备 什么时候使用? 慢性肺病患者，他们的通气依赖于缺氧驱动。 长时间高浓度供氧后，缺氧驱动会消失 通气不足可能发生，二氧化碳水平会变得非常高，导致昏迷甚至死亡。 |
| 100 | Title: Fixed Concentration Devices | 固定氧浓度设备 |
| 101 | Title: Fixed Concentration Devices Inspired O2 concentration is determined by the design of the device and NOT the patient’s rate and depth of breathing. | 固定氧浓度设备 吸入的氧浓度是由设备的设计，而不是病人的呼吸频率和深度决定的。 |
| 102 | Title: Fixed Concentration Devices Features High gas flow to the patient Venturi Effect | 固定氧浓度设备 特征 高流量气体输送给患者 文丘里效应 |
| 103 | Title: Fixed Venturi Flow rates Oxygen flow meter - litres per minute (L/min) Oxygen Concentration Total flow rate to patient Litres per minute (L/min) Note: Figures in white indicate total gas flows which may be less than the patient's peak inspiratory flow. Figures in Bold indicate total flow rate to patient when using recommended oxygen flow. | 固定文氏阀 流量 氧气流量计-L/分钟 氧气浓度 输送给病人的总流量 L/分钟 注意: 白色的数字表示总气体流量，可能小于患者的吸气峰值流量。在使用推荐的氧气流量时，用粗体表示病人的总流量。 |
| 104 | Title: Fixed Venturi Recommended flow for a 35% venturi valve is 8L/min. Caution 8L/min O2 +37 L/min Air = Total flow rate 45 L/min. Caution Increasing the oxygen flow rate does NOT increase the oxygen concentration. 10 L/min O2 +46L/min Air = Total flow rate 56 L/min- | 固定文氏阀 35%氧浓度的文氏阀的推荐流量为8L/分钟 注意 8L/分钟 氧气+37 L/分钟空气= 总流量45 L/分钟. 注意 提高氧气流量并不能提高氧气浓度 10L/分钟 氧气+ 46/分钟空气=总流量 56L/分钟 |
| 105 | Title: Oxygen Recovery T-Piece Kits Features Provides spontaneously breathing, intubated patients with supplemetary oxygen during transport and in the recovery room. | 氧气康复 T型组件 特征 用来为自主呼吸的插管患者在转移过程中或恢复室提供氧气 |
| 106 | Title: Oxygen Recovery T-Piece Kits Method of Use | 氧气康复 T型组件 使用方法 |
| 107 | Title: Adjustable Venturi Features Between 26% & 50% oxygen Twist & lock | 可调文氏阀 特征 26%-50%氧浓度 旋转&锁紧 |
| 108 | Title: Adjustable Venturi Method of use | 可调文氏阀 使用方法 |
| 109 | Title: Intersurgical EcoLite Oxygen Masks Environmental impact Using the simapro 67 life cycle assessment. Eco milli points score | 英特赛克Ecolite氧气面罩 环境影响 使用simapro 67生命周期评估 生态分数 |
| 110 | Title: Humidification Devices | 加湿设备 |
| 111 | Title: Bubble Humidifier Flow rate To the patient | 气泡加湿器 输送给病人的流量 |
| 112 | Title: Bubble Humidifier Humidity of the gas varies with: Flow rate Water temperature Bottle size Water Volume | 气泡加湿器 气体湿度会发生变化： 流量 水温 水瓶尺寸 水量 |
| 113 | Title: Bubble Humidifier Features - Commonly use with low flow devices - Pressure relief valve - Universal Screw head - Can be used with some prefilled sterile water bottles | 气泡加湿器 特征 一般用于低流量设备 压力释放阀门 通用螺纹头 可用于预罐装的无菌水瓶 |
| 114 | Title: Bubble Humidifier Method of use | 气泡加湿器 使用方法 |
| 115 | Title: Humidifier Nebuliser Oxygen Entrained air To patient Water | 气泡加湿器 氧气 吸入空气 病人端 水 |
| 116 | Title: Humidifier Nebuliser Features - Deliver 24-50mg H2O per litre of gas - Most commonly used with aerosol and tracheaostomy masks. | 气泡加湿器 特征 每升气体可提供24-50mg 水 广泛用于雾化面罩和气切面罩 |
| 117 | Title: Humidifier Nebuliser Features - Should Operate as quietly as possible - Should conform to international standards to ensure accurate O2 concentrations | 气泡加湿器 特征 尽可能安静地操作 应符合国际标准，以确保准确的氧气浓度 |
| 118 | Title: Humidifier Nebuliser Features - Some are tamperproof - Head may incorporate a Universal Screw Thread to take standard bottles | 气泡加湿器 特征 有些带有安全盖 瓶口有国际通用螺纹，适配标准瓶子 |
| 119 | Title: Humidifier Nebuliser Method of Use Caution The lowest point of the flextube must be lower than the patient to allow condensation to drain safely. | 气泡加湿器 使用方法 注意 伸缩管的最低点必须低于病人，这样冷凝水才能安全排出。 |
| 120 | Title: Humidifier Nebuliser | 气泡加湿器 |
| 121 | Title: Humidification Devices | 加湿设备 |
| 122 | Title: Humidification Devices Caution - Not for use in ventilator circuits - Sterile fluids only | 加湿设备 注意 不能用于呼吸回路 只能用无菌水 |
| 123 | Key Oxygen Exhalation Ambient air ETCO2 monitoring line | 关键 氧气 呼气 周围空气 呼吸末二氧化碳监测管 |
| 124 | Covered Variable oxygen concentration devices Fixed concenrtation devices Humidification devices | 囊括了： 可变氧浓度设备 固定氧浓度设备 加湿设备 |

References

AUDIT UPDATE, British Thoracic Society emergency oxygen audits, B R O’Driscoll,1 L S Howard,2 C Bucknall,3 S A Welham,4 A G Davison,5 On behalf of the British Thoracic Society Thorax 2011;66:734e735

British Thoracic Society Guidelines for the emergency use of oxygen in adult patients. Thorax 2008, Volume 63, Supplement VI

American Association for Respiratory Care Guideline: Oxygen Therapy for adults in the acute care facility, Respiratory Care June 2002 Vol 47 no. 6.

Clinical Practice in Respiratory Care, James B. Fink, Gerald E. Hunt, Lippincott Williams & Wilkins, 1999